## Sheet1
orginContent | matchContent | matchContentPercent | matchPage | matchRow | matchTitle | titleIsExist
Description | Description | 100% | 322 | 1 | Description | FALSE
UPS must be available. | t be available | 58.33% | 326 | 175 | UPS must be available. | FALSE
No measure required | e required | 52.63% | 226 | 172 | No measure required | FALSE
Check emergency stop during periodic maintenance. | periodic maintenance | 41.18% | 9 | 904 | Check emergency stop during periodic maintenance. | FALSE
CAUTION Cause: Loss of image data Possible consequences: Risk of unnecessary X-ray exposure Remedy: Do not shut down the system while radiation is being released. | Do not shut down the system while radiation is being released. | 38.18% | 89 | 1066 | CAUTION Cause: Loss of image data Possible consequences: Risk of unnecessary X-ray exposure Remedy: Do not shut down the system while radiation is being released. | TRUE
Mechanical end stop at the end of the movement range | at the end of the | 36.54% | 208 | 770 | Mechanical end stop at the end of the movement range | FALSE
CAUTION Cause: X-ray release may be possible from remote locations. Risk: Unnecessary dose. Remedy: Do not release X-ray mistakenly if not in viewing distance to the patient. | Do not release X-ray mistakenly if not in viewing distance to the | 36.87% | 304 | 345 | CAUTION Cause: X-ray release may be possible from remote locations. Risk: Unnecessary dose. Remedy: Do not release X-ray mistakenly if not in viewing distance to the patient. | TRUE
Automatic calibration procedure has to be initiated manually. | calibration procedure | 36.07% | 190 | 1031 | Automatic calibration procedure has to be initiated manually. | FALSE
CAUTION Cause: Improper cleaning Possible consequences: Risk of infection! Remedy: After dealing with patients with highly infectious diseases, clean the unit wheels, in addition to regular cleaning. | After dealing with patients with highly infectious diseases, clean the | 35.15% | 253 | 164 | CAUTION Cause: Improper cleaning Possible consequences: Risk of infection! Remedy: After dealing with patients with highly infectious diseases, clean the unit wheels, in addition to regular cleaning. | TRUE
Use of reliable data transfer standards and protocols. | able data transfer | 35.19% | 219 | 573 | Use of reliable data transfer standards and protocols. | FALSE
Collimator must be developed in accordance with the IEC 60601-1 standards. | in accordance with the IEC | 35.53% | 234 | 923 | Collimator must be developed in accordance with the IEC 60601-1 standards. | FALSE
The system must be tested for biocompatibility according to ISO 10993-1. | ompatibility according to | 35.14% | 49 | 939 | The system must be tested for biocompatibility according to ISO 10993-1. | FALSE
Provide training for technician | technician | 33.33% | 251 | 999 | Provide training for technician | FALSE
CAUTION Cause: Improper use of the system beyond its intended use. Possible consequences: Risk to the patient and user! Remedy: The system may only be used as defined by its intended use for clinical applications. | The system may only be used as defined by its intended use for clinical | 33.33% | 16 | 1231 | CAUTION Cause: Improper use of the system beyond its intended use. Possible consequences: Risk to the patient and user! Remedy: The system may only be used as defined by its intended use for clinical applications. | FALSE
Request for maintenance of the dose area product meter on a regular base. | of the dose area product | 33.33% | 141 | 364 | Request for maintenance of the dose area product meter on a regular base. | FALSE
A Virus Protection method is provided in the system. | s provided in the | 33.33% | 273 | 383 | A Virus Protection method is provided in the system. | FALSE
CAUTION Cause: Operating error Possible Consequences: Unnecessary radiation Remedy: Follow instruction of calibration procedure. | Follow instruction of calibration procedure. | 33.83% | 190 | 1033 | CAUTION Cause: Operating error Possible Consequences: Unnecessary radiation Remedy: Follow instruction of calibration procedure. | TRUE
Only images of the selected patient can be transferred to the hardcopy camera | of the selected patient | 32.47% | 167 | 548 | Only images of the selected patient can be transferred to the hardcopy camera | FALSE
Warning Cause: power supply network not be grounded. Possible consequence: Electric shock to operator and patient Remedy: Connect this device only to a power supply network that has been properly grounded. | Connect this device only to a power supply network that has been | 30.95% | 71 | 1101 | Warning Cause: power supply network not be grounded. Possible consequence: Electric shock to operator and patient Remedy: Connect this device only to a power supply network that has been properly grounded. | TRUE
Acquired images are always stored immediately on hard disk. | quired images are | 30.51% | 175 | 651 | Acquired images are always stored immediately on hard disk. | FALSE
Warning Cause: System component failure during cardiac/vascular procedure. Possible consequences: Risk of injury for the patient!Remedy: Prepare a reserve system during cardiac/vascular applications. | Prepare a reserve system during cardiac/vascular applications. | 31.34% | 27 | 997 | Warning Cause: System component failure during cardiac/vascular procedure. Possible consequences: Risk of injury for the patient!Remedy: Prepare a reserve system during cardiac/vascular applications. | TRUE
IQAP measurement after each dose adjustment. | measurement a | 31.82% | 193 | 532 | IQAP measurement after each dose adjustment. | FALSE
no measure required due to risk classification | classification | 30.61% | 281 | 646 | no measure required due to risk classification | FALSE
CAUTION Cause: Distance between patient or other persons and system is too small. Possible consequences: Risk of crushing! Remedy: Use the system only if the patient and other persons can be observed during movements. | Use the system only if the patient and other persons can be observed | 31.08% | 74 | 753 | CAUTION Cause: Distance between patient or other persons and system is too small. Possible consequences: Risk of crushing! Remedy: Use the system only if the patient and other persons can be observed during movements. | TRUE
no measure required due to risk classification | classification | 30.61% | 281 | 613 | no measure required due to risk classification | FALSE
no measure required due to risk classification | classification | 30.61% | 281 | 610 | no measure required due to risk classification | FALSE
no measure required due to risk classification | classification | 30.61% | 281 | 629 | no measure required due to risk classification | FALSE
no measure required due to risk classification | classification | 30.61% | 281 | 620 | no measure required due to risk classification | FALSE
CAUTION Cause: Distance between patient or other persons and system is too small. Possible consequences: Risk of crushing! Remedy: Use the system only if the patient and other persons can be observed during movements. | Use the system only if the patient and other persons can be observed | 31.36% | 74 | 756 | CAUTION Cause: Distance between patient or other persons and system is too small. Possible consequences: Risk of crushing! Remedy: Use the system only if the patient and other persons can be observed during movements. | TRUE
Warning Cause: Increase of tilt moment Possible consequences: Risk of crushing Remedy: Only products/components approved by Siemens may be attached. | Only products/components approved by Siemens | 30.46% | 95 | 1270 | Warning Cause: Increase of tilt moment Possible consequences: Risk of crushing Remedy: Only products/components approved by Siemens may be attached. | TRUE
Attention note in the relevant service manual: There is a risk of eye injury! Do not look directly into laser-beam. Do not observe directly with optical instruments. | Do not observe directly with optical instruments. | 29.17% | 309 | 204 | Attention note in the relevant service manual: There is a risk of eye injury! Do not look directly into laser-beam. Do not observe directly with optical instruments. | FALSE
CAUTIONCause: Put the liquid on the system. Possible consequences: Risk of electrical hazard or damage to the system Remedy: Do not put the liquid on the system. | Risk of electrical hazard or damage to the system | 30.06% | 256 | 945 | CAUTIONCause: Put the liquid on the system. Possible consequences: Risk of electrical hazard or damage to the system Remedy: Do not put the liquid on the system. | TRUE
CAUTION Cause: Inverted image orientation. Possible consequences: Risk of an incorrect decision during surgery. Remedy: Prior to releasing radiation, the user must ensure the correct patient position with respect to the X-ray beam. | Prior to releasing radiation, the user must ensure the correct patient | 30.08% | 138 | 685 | CAUTION Cause: Inverted image orientation. Possible consequences: Risk of an incorrect decision during surgery. Remedy: Prior to releasing radiation, the user must ensure the correct patient position with respect to the X-ray beam. | TRUE
Fuses shall be provided in each supply lead | provided in | 28.89% | 273 | 1098 | Fuses shall be provided in each supply lead | FALSE
see EPH System Architecture: | System | 26.67% | 315 | 1010 | see EPH System Architecture: | FALSE
A message is displayed to restart the system when the Service environment is left. | A message is displayed | 27.38% | 242 | 82 | A message is displayed to restart the system when the Service environment is left. | FALSE
Warning Cause: Change of over-all center of gravity Possible consequences: Risk of crushing Remedy: Only products/components approved by Siemens may be attached. | Only products/components approved by Siemens | 28.05% | 95 | 1166 | Warning Cause: Change of over-all center of gravity Possible consequences: Risk of crushing Remedy: Only products/components approved by Siemens may be attached. | TRUE
Do not connect any additional device to the UPS power outlets | additional device | 28.57% | 28 | 245 | Do not connect any additional device to the UPS power outlets | FALSE
Protective earth measurement during final inspection | rotective earth | 28.85% | 71 | 1193 | Protective earth measurement during final inspection | FALSE
Warning Cause: Interference Possible consequences: Effect on electronic life-supporting device on patient Remedy: When observing uncommon system behavior (performance characteristics) additional measures (adjustment, relocation) could be necessary. | characteristics) additional measures (adjustment, relocation) could be | 27.67% | 49 | 1173 | Warning Cause: Interference Possible consequences: Effect on electronic life-supporting device on patient Remedy: When observing uncommon system behavior (performance characteristics) additional measures (adjustment, relocation) could be necessary. | FALSE
CAUTION Cause: time gap between display of K-wire including overlay grafic and real moved K-wire Possible consequences: Risk of wrong interventional decision Remedy: Duty of care for operator, that he/she shall be aware of this possible mismatch | Duty of care for operator, that he/she shall be aware of this possible | 28.63% | 156 | 660 | CAUTION Cause: time gap between display of K-wire including overlay grafic and real moved K-wire Possible consequences: Risk of wrong interventional decision Remedy: Duty of care for operator, that he/she shall be aware of this possible mismatch | TRUE
The selected clinical procedure shall be shown on the monitor during examination. | e shown on the monitor | 28.4% | 119 | 526 | The selected clinical procedure shall be shown on the monitor during examination. | FALSE
AT-S3-28911 - Missing cross-reference is sufficient. | e is sufficient | 28.85% | 217 | 717 | AT-S3-28911 - Missing cross-reference is sufficient. | FALSE
Covers are designed that openings and gaps are reduced. | s are designed | 26.32% | 43 | 483 | Covers are designed that openings and gaps are reduced. | FALSE
Radiation function is forbidden during standby mode. | ation function | 26.92% | 189 | 388 | Radiation function is forbidden during standby mode. | FALSE
CAUTION Cause: Network connection may not be available continuously. Possible consequences: Reference images not available during intervention. Remedy: Before beginning the intervention, the user should load the required images from the network archive. | Before beginning the intervention, the user should load the required | 26.74% | 106 | 87 | CAUTION Cause: Network connection may not be available continuously. Possible consequences: Reference images not available during intervention. Remedy: Before beginning the intervention, the user should load the required images from the network archive. | TRUE
Warning: Cause: Connecting the LAN socket on the trolley in operation room Possible consequences: Risk of electric shock Remedy: Only IT equipments/products approved by IEC 60950-1 or IEC 60601-1 can be used in connection with this system. | Only IT equipments/products approved by IEC 60950-1 or IEC 60601-1 | 27.69% | 48 | 1093 | Warning: Cause: Connecting the LAN socket on the trolley in operation room Possible consequences: Risk of electric shock Remedy: Only IT equipments/products approved by IEC 60950-1 or IEC 60601-1 can be used in connection with this system. | TRUE
No direct contact possible due to access cover of the detector. | of the detector. | 26.98% | 276 | 715 | No direct contact possible due to access cover of the detector. | FALSE
Leakage current measurement during final inspection | measurement d | 27.45% | 247 | 1197 | Leakage current measurement during final inspection | FALSE
When the user starts radiation, imaging system clears the LIH image. | imaging system c | 25% | 91 | 664 | When the user starts radiation, imaging system clears the LIH image. | FALSE
The system shall indicate the battery charge for standby mode. | battery charge | 25.81% | 61 | 389 | The system shall indicate the battery charge for standby mode. | FALSE
The autorun feature for removable media must be disabled on the imaging system. | n the imaging system. | 25.93% | 326 | 1059 | The autorun feature for removable media must be disabled on the imaging system. | FALSE
CAUTION Cause: Interference Possible consequences: Effect on people with implant Remedy: A minimum distance has to be kept according to IEC 60601-1-2. If this not possible, the WLAN-client is to be turned off. | this not possible, the WLAN-client is to be turned off. | 25.94% | 300 | 1016 | CAUTION Cause: Interference Possible consequences: Effect on people with implant Remedy: A minimum distance has to be kept according to IEC 60601-1-2. If this not possible, the WLAN-client is to be turned off. | FALSE
A label shall be attached on wireless footswitch to which system he belongs to. | wireless footswitch | 24.71% | 328 | 350 | A label shall be attached on wireless footswitch to which system he belongs to. | FALSE
CARION Cause: Software lockup Possible consequences: Work with the imaging system is not possible anymore. Remedy: Restart the system.. | the imaging system is not possible | 25% | 325 | 85 | CARION Cause: Software lockup Possible consequences: Work with the imaging system is not possible anymore. Remedy: Restart the system.. | TRUE
Capacitors are mounted under covers. | ed under | 25% | 273 | 1087 | Capacitors are mounted under covers. | FALSE
Adjustment of the collimator plate is designed so that the edge is visible in the fluoro image. | is visible in the fluoro | 25.77% | 131 | 501 | Adjustment of the collimator plate is designed so that the edge is visible in the fluoro image. | FALSE
It has to be detected at the system if grid is inserted or not | at the system i | 23.44% | 248 | 488 | It has to be detected at the system if grid is inserted or not | FALSE
There should be at least 2 indicator lamps in a system. | be at least | 23.64% | 192 | 1073 | There should be at least 2 indicator lamps in a system. | FALSE
AT-S3-28454 - Missing cross-reference | reference | 24.32% | 325 | 696 | AT-S3-28454 - Missing cross-reference | FALSE
A regular maintenance is recommended to ensure always a best possible of image quality. | is recommended to en | 24.14% | 300 | 601 | A regular maintenance is recommended to ensure always a best possible of image quality. | FALSE
If the imaging system or detector is not ready for radiation, the system blocks any radiation release request. | is not ready for radiation | 24.55% | 73 | 527 | If the imaging system or detector is not ready for radiation, the system blocks any radiation release request. | FALSE
Storage of the image number in the image header of the 2D images. | ge of the image | 24.62% | 194 | 461 | Storage of the image number in the image header of the 2D images. | FALSE
The system shall show patient name on monitor A during examination. | ing examination. | 23.88% | 147 | 464 | The system shall show patient name on monitor A during examination. | FALSE
CAUTION Cause: Radiation released from a footswitch in another room Risk: Non-clinical exposure to ionizing radiation Remedy: As part of the system installation, the wireless footswitch shall be labeled according to which system it belongs to. | As part of the system installation, the wireless footswitch | 24.6% | 303 | 352 | CAUTION Cause: Radiation released from a footswitch in another room Risk: Non-clinical exposure to ionizing radiation Remedy: As part of the system installation, the wireless footswitch shall be labeled according to which system it belongs to. | TRUE
AT-S3-29121 - Missing cross-reference | reference | 24.32% | 325 | 487 | AT-S3-29121 - Missing cross-reference | FALSE
CAUTION Cause: Risk that CD/DVD-R media used for exporting can become damaged or be unreadable on other CD/DVD drives. Possible consequences: Data loss or apparent data loss. Remedy: Use removable devices as transfer media only. Only data media approved by Siemens may be used. | Risk that CD/DVD-R media used for exporting can become damaged or be | 23.94% | 208 | 74 | CAUTION Cause: Risk that CD/DVD-R media used for exporting can become damaged or be unreadable on other CD/DVD drives. Possible consequences: Data loss or apparent data loss. Remedy: Use removable devices as transfer media only. Only data media approved by Siemens may be used. | TRUE
The radiation power of the WLAN components has to be compliant with IEC 60601-1-2. | with IEC 60601-1-2. | 24.39% | 300 | 1021 | The radiation power of the WLAN components has to be compliant with IEC 60601-1-2. | FALSE
Place warning label (pinch label) next to foot brake of trolley. | warning label ( | 25% | 292 | 1243 | Place warning label (pinch label) next to foot brake of trolley. | FALSE
Compliance with IEC 60601-1 chap. 9.2 Table 20, Acceptable gaps for finger > 25mm | ance with IEC 60601-1 | 25.3% | 300 | 758 | Compliance with IEC 60601-1 chap. 9.2 Table 20, Acceptable gaps for finger > 25mm | FALSE
CAUTION Cause: overlay itself may hide K-wire in live X-ray image (depends on thickness of K-Wire and overlay configuration) Possible consequences: Risk of wrong interventional decision Remedy: K-wire overlay feature can be switched-off (toggle ON/OFF) to allow free view on X-ray live image | verlay itself may hide K-wire in live X-ray image (depends on thickness of | 25.17% | 156 | 658 | CAUTION Cause: overlay itself may hide K-wire in live X-ray image (depends on thickness of K-Wire and overlay configuration) Possible consequences: Risk of wrong interventional decision Remedy: K-wire overlay feature can be switched-off (toggle ON/OFF) to allow free view on X-ray live image | TRUE
Check emergency stop during periodic maintenance. See AT-S3-28727 - Missing cross-reference | periodic maintenance | 22.83% | 9 | 890 | Check emergency stop during periodic maintenance. See AT-S3-28727 - Missing cross-reference | FALSE
Customer has to perform regular maintenance Same as AT-S3-28630 - Missing cross-reference | regular maintenance | 22.22% | 23 | 1069 | Customer has to perform regular maintenance Same as AT-S3-28630 - Missing cross-reference | FALSE
CAUTION Cause: C-arm movements during operations Possible consequences: Risk of injuries Remedy: Due to the maneuverability of the C-arm, collisions with other OR-instruments are possible if the device is not used according to the intended use and instructions provided in this manual. | Due to the maneuverability of the C-arm, collisions with other | 22.22% | 111 | 747 | CAUTION Cause: C-arm movements during operations Possible consequences: Risk of injuries Remedy: Due to the maneuverability of the C-arm, collisions with other OR-instruments are possible if the device is not used according to the intended use and instructions provided in this manual. | TRUE
Caution: Cause: Manual or motorized movements Possible consequences: Risks of injury by crushing or collision for persons and objects in the vicinity of moving parts. Remedy: check functionality of emergency stop on a daily basis | collision for persons and objects in the vicinity of | 22.84% | 74 | 905 | Caution: Cause: Manual or motorized movements Possible consequences: Risks of injury by crushing or collision for persons and objects in the vicinity of moving parts. Remedy: check functionality of emergency stop on a daily basis | FALSE
Product specific service training (for service calibration) | ce calibration | 23.73% | 192 | 1028 | Product specific service training (for service calibration) | FALSE
When patient or image data has been corrected, the imaging system refreshes the display. | , the imaging system | 23.86% | 61 | 550 | When patient or image data has been corrected, the imaging system refreshes the display. | FALSE
Label:forbideen to touch the metal parts of the screw stem | parts of the | 24.14% | 229 | 970 | Label:forbideen to touch the metal parts of the screw stem | FALSE
The system shall indicate the tube loading factors at the monitor. | at the monitor | 22.73% | 301 | 560 | The system shall indicate the tube loading factors at the monitor. | FALSE
Cable holders must be developed in accordance with the IEC 60601-1 standards. AT-S3-28411 - Missing cross-reference | in accordance with the IEC | 22.5% | 234 | 925 | Cable holders must be developed in accordance with the IEC 60601-1 standards. AT-S3-28411 - Missing cross-reference | FALSE
Place a warning label (No pushing, ISO 7010-P017) on monitor holder. | warning label ( | 22.86% | 292 | 1269 | Place a warning label (No pushing, ISO 7010-P017) on monitor holder. | FALSE
AT-S3-29035 - Missing cross-reference | reference | 23.08% | 325 | 595 | AT-S3-29035 - Missing cross-reference | FALSE
Encapsulation and separate software tests for brightness value calculation. | brightness value | 22.67% | 184 | 525 | Encapsulation and separate software tests for brightness value calculation. | FALSE
Ensure location of crushing zone switch during final inspection. | switch during | 21.21% | 304 | 796 | Ensure location of crushing zone switch during final inspection. | FALSE
An initial pairing has to be done. This pairing shall remain, even in the case of reboot, reconnection of the wireless footswitch or the system. | tion of the wireless footswitch | 21.53% | 303 | 351 | An initial pairing has to be done. This pairing shall remain, even in the case of reboot, reconnection of the wireless footswitch or the system. | FALSE
see AT-S3-29472 - Missing cross-reference | reference | 21.95% | 325 | 162 | see AT-S3-29472 - Missing cross-reference | FALSE
Caution in the user manual: Cause: Software lockup Possible consequences: Work with the imaging system is not possible anymore. Remedy: Restart the system.. | the imaging system is not possible | 21.74% | 325 | 48 | Caution in the user manual: Cause: Software lockup Possible consequences: Work with the imaging system is not possible anymore. Remedy: Restart the system.. | FALSE
A technical instruction for the disposal procedure is existing. | tion for the d | 21.54% | 264 | 936 | A technical instruction for the disposal procedure is existing. | FALSE
CAUTION Cause: Looking into laser-beam Possible consequences: Risk of impaired vision Remedy: There is a risk of eye injury! Do not look directly into laser-beam. Do not observe directly with optical instruments. | Do not observe directly with optical instruments. | 22.48% | 309 | 848 | CAUTION Cause: Looking into laser-beam Possible consequences: Risk of impaired vision Remedy: There is a risk of eye injury! Do not look directly into laser-beam. Do not observe directly with optical instruments. | TRUE
Note in service documentation: Correct FD calibration in regular intervals | regular intervals | 23.38% | 124 | 67 | Note in service documentation: Correct FD calibration in regular intervals | FALSE
Control elements must not stick out, but have to be designed flush or sunk in. | Control elements | 21.25% | 56 | 811 | Control elements must not stick out, but have to be designed flush or sunk in. | FALSE
see AT-S3-28600 - Missing cross-reference | reference | 21.95% | 325 | 431 | see AT-S3-28600 - Missing cross-reference | FALSE
see AT-S3-28603 - Missing cross-reference | reference | 21.95% | 325 | 426 | see AT-S3-28603 - Missing cross-reference | FALSE
System status is visible to indicate the presence of an error. | indicate the | 22.58% | 218 | 1064 | System status is visible to indicate the presence of an error. | FALSE
Same as AT-S3-28437 - Missing cross-reference | reference | 20% | 325 | 991 | Same as AT-S3-28437 - Missing cross-reference | FALSE
Same as AT-S3-29060 - Missing cross-reference | reference | 20% | 325 | 563 | Same as AT-S3-29060 - Missing cross-reference | FALSE
Same as AT-S3-29078 - Missing cross-reference | reference | 20% | 325 | 562 | Same as AT-S3-29078 - Missing cross-reference | FALSE
same as AT-S3-28402 - Missing cross-reference | reference | 20% | 325 | 994 | same as AT-S3-28402 - Missing cross-reference | FALSE
same as AT-S3-29078 - Missing cross-reference | reference | 20% | 325 | 566 | same as AT-S3-29078 - Missing cross-reference | FALSE
Same as AT-S3-28437 - Missing cross-reference | reference | 20% | 325 | 982 | Same as AT-S3-28437 - Missing cross-reference | FALSE
same asAT-S3-28363 - Missing cross-reference | reference | 20.45% | 325 | 1258 | same asAT-S3-28363 - Missing cross-reference | FALSE
Same as AT-S3-28437 - Missing cross-reference | reference | 20% | 325 | 988 | Same as AT-S3-28437 - Missing cross-reference | FALSE
same as AT-S3-28397 - Missing cross-reference | reference | 20% | 325 | 1003 | same as AT-S3-28397 - Missing cross-reference | FALSE
same asAT-S3-28360 - Missing cross-reference | reference | 20.45% | 325 | 1260 | same asAT-S3-28360 - Missing cross-reference | FALSE
Same as AT-S3-28600 - Missing cross-reference | reference | 20% | 325 | 1025 | Same as AT-S3-28600 - Missing cross-reference | FALSE
Same as AT-S3-28402 - Missing cross-reference | reference | 20% | 325 | 1005 | Same as AT-S3-28402 - Missing cross-reference | FALSE
Same as AT-S3-29006 - Missing cross-reference | reference | 20% | 325 | 607 | Same as AT-S3-29006 - Missing cross-reference | FALSE
same as AT-S3-28417 - Missing cross-reference | reference | 20% | 325 | 942 | same as AT-S3-28417 - Missing cross-reference | FALSE
Same as AT-S3-28411 - Missing cross-reference | reference | 20% | 325 | 975 | Same as AT-S3-28411 - Missing cross-reference | FALSE
Same as AT-S3-28790 - Missing cross-reference | reference | 20% | 325 | 900 | Same as AT-S3-28790 - Missing cross-reference | FALSE
Same as AT-S3-28789 - Missing cross-reference | reference | 20% | 325 | 901 | Same as AT-S3-28789 - Missing cross-reference | FALSE
Same as AT-S3-28824 - Missing cross-reference | reference | 20% | 325 | 739 | Same as AT-S3-28824 - Missing cross-reference | FALSE
Same as AT-S3-29280 - Missing cross-reference | reference | 20% | 325 | 883 | Same as AT-S3-29280 - Missing cross-reference | FALSE
Same as AT-S3-28551 - Missing cross-reference | reference | 20% | 325 | 841 | Same as AT-S3-28551 - Missing cross-reference | FALSE
Same as AT-S3-28790 - Missing cross-reference | reference | 20% | 325 | 837 | Same as AT-S3-28790 - Missing cross-reference | FALSE
Same as AT-S3-28551 - Missing cross-reference | reference | 20% | 325 | 838 | Same as AT-S3-28551 - Missing cross-reference | FALSE
same as AT-S3-28438 - Missing cross-reference | reference | 20% | 325 | 832 | same as AT-S3-28438 - Missing cross-reference | FALSE
same as AT-S3-28438 - Missing cross-reference | reference | 20% | 325 | 829 | same as AT-S3-28438 - Missing cross-reference | FALSE
Specify sufficient braking power. This is in compliance with IEC 60601-1; clause 9.4.3 (Instability). | ance with IEC 60601-1 | 20.59% | 300 | 763 | Specify sufficient braking power. This is in compliance with IEC 60601-1; clause 9.4.3 (Instability). | FALSE
same as AT-S3-28692 - Missing cross-reference | reference | 20% | 325 | 821 | same as AT-S3-28692 - Missing cross-reference | FALSE
Same as AT-S3-28665 - Missing cross-reference | reference | 20.83% | 310 | 818 | Same as AT-S3-28665 - Missing cross-reference | FALSE
Same as AT-S3-28809 - Missing cross-reference | reference | 20% | 325 | 815 | Same as AT-S3-28809 - Missing cross-reference | FALSE
same as AT-S3-28789 - Missing cross-reference | reference | 20% | 325 | 793 | same as AT-S3-28789 - Missing cross-reference | FALSE
same as AT-S3-29158 - Missing cross-reference | reference | 20% | 325 | 27 | same as AT-S3-29158 - Missing cross-reference | FALSE
same as AT-S3-28914 - Missing cross-reference | reference | 20% | 325 | 726 | same as AT-S3-28914 - Missing cross-reference | FALSE
Same as AT-S3-28824 - Missing cross-reference | reference | 20% | 325 | 912 | Same as AT-S3-28824 - Missing cross-reference | FALSE
Same as AT-S3-28908 - Missing cross-reference | reference | 20% | 325 | 722 | Same as AT-S3-28908 - Missing cross-reference | FALSE
The monoblock design shall fulfill requirement of IEC 60601-2-28 Chap. 9.7. | IEC 60601-2-28 | 19.48% | 323 | 712 | The monoblock design shall fulfill requirement of IEC 60601-2-28 Chap. 9.7. | FALSE
Compliance with DIN EN 60601-1-3 (2013); Chap. 11 residual radiation. | ance with DIN | 20.29% | 277 | 640 | Compliance with DIN EN 60601-1-3 (2013); Chap. 11 residual radiation. | FALSE
Same as AT-S3-28411 - Missing cross-reference | reference | 20% | 325 | 959 | Same as AT-S3-28411 - Missing cross-reference | FALSE
same as AT-S3-28958 - Missing cross-reference | reference | 20% | 325 | 673 | same as AT-S3-28958 - Missing cross-reference | TRUE
same as AT-S3-28948 - Missing cross-reference | reference | 20% | 325 | 680 | same as AT-S3-28948 - Missing cross-reference | FALSE
Same as AT-S3-28943 - Missing cross-reference | reference | 20% | 325 | 946 | Same as AT-S3-28943 - Missing cross-reference | FALSE
Same as AT-S3-28943 - Missing cross-reference | reference | 20% | 325 | 949 | Same as AT-S3-28943 - Missing cross-reference | FALSE
Housing of monoblock prevents an intrusion of liquids. | of liquids | 19.64% | 307 | 689 | Housing of monoblock prevents an intrusion of liquids. | FALSE
Same as AT-S3-28411 - Missing cross-reference | reference | 20% | 325 | 978 | Same as AT-S3-28411 - Missing cross-reference | FALSE
The imaging system software triggers the start of the mask calculation phase depending on the system state (subtraction or roadmap). | he imaging system software | 20.15% | 67 | 671 | The imaging system software triggers the start of the mask calculation phase depending on the system state (subtraction or roadmap). | FALSE
Same as AT-S3-28551 - Missing cross-reference | reference | 20% | 325 | 920 | Same as AT-S3-28551 - Missing cross-reference | FALSE
same as AT-S3-28921 - Missing cross-reference | reference | 20% | 325 | 705 | same as AT-S3-28921 - Missing cross-reference | FALSE
same as AT-S3-28908 - Missing cross-reference | reference | 20% | 325 | 710 | same as AT-S3-28908 - Missing cross-reference | FALSE
same as AT-S3-28914 - Missing cross-reference | reference | 20% | 325 | 711 | same as AT-S3-28914 - Missing cross-reference | FALSE
same as AT-S3-28921 - Missing cross-reference | reference | 20% | 325 | 707 | same as AT-S3-28921 - Missing cross-reference | FALSE
Same as AT-S3-28713 - Missing cross-reference | reference | 20% | 325 | 914 | Same as AT-S3-28713 - Missing cross-reference | FALSE
Same as AT-S3-28714 - Missing cross-reference | reference | 20% | 325 | 913 | Same as AT-S3-28714 - Missing cross-reference | FALSE
same as AT-S3-28423 - Missing cross-reference | reference | 20% | 325 | 933 | same as AT-S3-28423 - Missing cross-reference | FALSE
same as AT-S3-29480 - Missing cross-reference | reference | 20% | 325 | 174 | same as AT-S3-29480 - Missing cross-reference | FALSE
same asAT-S3-28365 - Missing cross-reference | reference | 20.45% | 325 | 1256 | same asAT-S3-28365 - Missing cross-reference | FALSE
Check that parameter values are within the defined range. | within the | 21.05% | 274 | 555 | Check that parameter values are within the defined range. | FALSE
Warning Cause: Damage of the cable Possible consequences: Risk of electric shock Remedy: Check power cable. If the power cable is impaired, the device must not be used. | power cable. If the power cable is | 21.05% | 242 | 1116 | Warning Cause: Damage of the cable Possible consequences: Risk of electric shock Remedy: Check power cable. If the power cable is impaired, the device must not be used. | TRUE
same as AT-S3-28571 - Missing cross-reference | reference | 20% | 325 | 282 | same as AT-S3-28571 - Missing cross-reference | FALSE
see AT-S3-28935 - Missing cross-reference | reference | 20.93% | 325 | 1178 | see AT-S3-28935 - Missing cross-reference | FALSE
CAUTION Cause: Incorrect image processing due to wrong configured system language Possible consequences: Incorrect basis for diagnosis Remedy: Configure the system only in one of the following languages: ......*1) *1) Enter only those system languages in operator manual which are validated by testcenter | Incorrect image processing due to wrong configured system language | 21.43% | 50 | 407 | CAUTION Cause: Incorrect image processing due to wrong configured system language Possible consequences: Incorrect basis for diagnosis Remedy: Configure the system only in one of the following languages: ......*1) *1) Enter only those system languages in operator manual which are validated by testcenter | TRUE
Utilize plug with stabilized mechanical connection against disconnection. | ical connection | 21.33% | 272 | 1076 | Utilize plug with stabilized mechanical connection against disconnection. | FALSE
Detailed monitor arm dismantling procedure shall be documented in service manual. | e documented in | 19.75% | 24 | 267 | Detailed monitor arm dismantling procedure shall be documented in service manual. | FALSE
Same as AT-S3-29251 - Missing cross-reference | reference | 20% | 325 | 280 | Same as AT-S3-29251 - Missing cross-reference | FALSE
same as AT-S3-28467 - Missing cross-reference | reference | 20% | 325 | 1155 | same as AT-S3-28467 - Missing cross-reference | FALSE
same as AT-S3-28377 - Missing cross-reference | reference | 20% | 325 | 1150 | same as AT-S3-28377 - Missing cross-reference | FALSE
same as AT-S3-28567 - Missing cross-reference | reference | 20% | 325 | 169 | same as AT-S3-28567 - Missing cross-reference | FALSE
same as AT-S3-28465 - Missing cross-reference | reference | 20% | 325 | 1157 | same as AT-S3-28465 - Missing cross-reference | FALSE
Plausibility and consistency check of worklist data. Worklistdata are marked with a flag. | are marked with a | 20.22% | 182 | 327 | Plausibility and consistency check of worklist data. Worklistdata are marked with a flag. | FALSE
same as AT-S3-28469 - Missing cross-reference | reference | 20% | 325 | 1153 | same as AT-S3-28469 - Missing cross-reference | FALSE
same as AT-S3-28380 - Missing cross-reference | reference | 20% | 325 | 1147 | same as AT-S3-28380 - Missing cross-reference | FALSE
Warning Labels show where voltage is present after shutdown. | e voltage is | 20.97% | 243 | 181 | Warning Labels show where voltage is present after shutdown. | FALSE
CAUTION Cause: Unauthorized or incorrect alterations/changes of the software or connection of the system to a network. Possible consequences: Risk of unauthorized access. Remedy: Make sure that all necessary precautions are taken with regard to the applicable security level if a functionality is added or the configuration of the condition on delivery is changed. | applicable security level if a functionality is added or the configuration | 20.16% | 17 | 1060 | CAUTION Cause: Unauthorized or incorrect alterations/changes of the software or connection of the system to a network. Possible consequences: Risk of unauthorized access. Remedy: Make sure that all necessary precautions are taken with regard to the applicable security level if a functionality is added or the configuration of the condition on delivery is changed. | TRUE
see AT-S3-28515 - Missing cross-reference | reference | 19.15% | 325 | 1108 | see AT-S3-28515 - Missing cross-reference | FALSE
same as AT-S3-28437 - Missing cross-reference | reference | 20% | 325 | 1114 | same as AT-S3-28437 - Missing cross-reference | FALSE
An intrusion detection solution based on whitelisting is provided by the system. | provided by the | 19.51% | 298 | 360 | An intrusion detection solution based on whitelisting is provided by the system. | FALSE
same as AT-S3-28437 - Missing cross-reference | reference | 20% | 325 | 1111 | same as AT-S3-28437 - Missing cross-reference | FALSE
Same as AT-S3-28630 - Missing cross-reference | reference | 19.15% | 325 | 1224 | Same as AT-S3-28630 - Missing cross-reference | FALSE
same as AT-S3-28439 - Missing cross-reference | reference | 20% | 325 | 1113 | same as AT-S3-28439 - Missing cross-reference | FALSE
Same as AT-S3-28397 - Missing cross-reference | reference | 20% | 325 | 1226 | Same as AT-S3-28397 - Missing cross-reference | FALSE
Same as AT-S3-28569 - Missing cross-reference | reference | 20% | 325 | 276 | Same as AT-S3-28569 - Missing cross-reference | FALSE
Specify sufficient braking power. Thi is Compliance with DIN EN 60601-1; Kap. 9.4.3(Instability) | sufficient braking | 19.79% | 328 | 1142 | Specify sufficient braking power. Thi is Compliance with DIN EN 60601-1; Kap. 9.4.3(Instability) | FALSE
Same as AT-S3-28722 - Missing cross-reference | reference | 19.15% | 325 | 908 | Same as AT-S3-28722 - Missing cross-reference | FALSE
Same as AT-S3-28679 - Missing cross-reference | reference | 20% | 325 | 261 | Same as AT-S3-28679 - Missing cross-reference | FALSE
Same as AT-S3-29480 - Missing cross-reference | reference | 20% | 325 | 177 | Same as AT-S3-29480 - Missing cross-reference | FALSE
same as AT-S3-28654 - Missing cross-reference | reference | 20% | 325 | 1217 | same as AT-S3-28654 - Missing cross-reference | FALSE
CAUTION Cause: Distance between operator rail and other components too small. Possible consequences: Risk of crushing! Remedy: Pay attention to the hazardous locations indicated. | the hazardous locations indicated. | 19.34% | 34 | 757 | CAUTION Cause: Distance between operator rail and other components too small. Possible consequences: Risk of crushing! Remedy: Pay attention to the hazardous locations indicated. | TRUE
same as AT-S3-29524 - Missing cross-reference | reference | 20% | 325 | 180 | same as AT-S3-29524 - Missing cross-reference | FALSE
same as AT-S3-28695 - Missing cross-reference | reference | 19.15% | 325 | 1210 | same as AT-S3-28695 - Missing cross-reference | FALSE
Same as AT-S3-28437 - Missing cross-reference | reference | 19.15% | 325 | 827 | Same as AT-S3-28437 - Missing cross-reference | FALSE
same as AT-S3-28603 - Missing cross-reference | reference | 20% | 325 | 186 | same as AT-S3-28603 - Missing cross-reference | FALSE
same as AT-S3-28943 - Missing cross-reference | reference | 19.15% | 325 | 26 | same as AT-S3-28943 - Missing cross-reference | FALSE
Same as AT-S3-28437 - Missing cross-reference | reference | 19.15% | 325 | 830 | Same as AT-S3-28437 - Missing cross-reference | FALSE
same as AT-S3-28463 - Missing cross-reference | reference | 20% | 325 | 1159 | same as AT-S3-28463 - Missing cross-reference | FALSE
same as AT-S3-28397 - Missing cross-reference | reference | 19.15% | 325 | 187 | same as AT-S3-28397 - Missing cross-reference | FALSE
same as AT-S3-28784 - Missing cross-reference | reference | 20% | 325 | 206 | same as AT-S3-28784 - Missing cross-reference | FALSE
Same as AT-S3-28402 - Missing cross-reference | reference | 19.15% | 325 | 1222 | Same as AT-S3-28402 - Missing cross-reference | TRUE
The main unit blocks automated movement if control elements are already activated during start-up. | f control elements | 19.39% | 53 | 845 | The main unit blocks automated movement if control elements are already activated during start-up. | FALSE
Work step in maintenance documents: Scanning of the conducting elements for damage and fouling | ting elements for | 19.15% | 84 | 1186 | Work step in maintenance documents: Scanning of the conducting elements for damage and fouling | FALSE
Use of reliable data transfer standards and protocols. Error messages in case of network problems. | able data transfer | 19.39% | 219 | 865 | Use of reliable data transfer standards and protocols. Error messages in case of network problems. | FALSE
same as AT-S3-28437 - Missing cross-reference | reference | 20% | 325 | 1183 | same as AT-S3-28437 - Missing cross-reference | FALSE
same as AT-S3-28439 - Missing cross-reference | reference | 20% | 325 | 1181 | same as AT-S3-28439 - Missing cross-reference | FALSE
Same as AT-S3-28683 - Missing cross-reference | reference | 19.15% | 325 | 886 | Same as AT-S3-28683 - Missing cross-reference | FALSE
same as AT-S3-28530 - Missing cross-reference | reference | 19.15% | 325 | 1104 | same as AT-S3-28530 - Missing cross-reference | FALSE
Same as AT-S3-28437 - Missing cross-reference | reference | 19.15% | 325 | 833 | Same as AT-S3-28437 - Missing cross-reference | FALSE
same as AT-S3-28363 - Missing cross-reference | reference | 19.15% | 325 | 1240 | same as AT-S3-28363 - Missing cross-reference | FALSE
same as AT-S3-28439 - Missing cross-reference | reference | 20% | 325 | 1110 | same as AT-S3-28439 - Missing cross-reference | FALSE
same as AT-S3-28397 - Missing cross-reference | reference | 20% | 325 | 1056 | same as AT-S3-28397 - Missing cross-reference | FALSE
see AT-S3-29111 - Missing cross-reference | reference | 20.93% | 325 | 518 | see AT-S3-29111 - Missing cross-reference | FALSE
see AT-S3-28943 - Missing cross-reference | reference | 20.93% | 325 | 478 | see AT-S3-28943 - Missing cross-reference | FALSE
Built-in parts have to be secured to avoid risk of dropping. | have to be s | 20.97% | 141 | 1237 | Built-in parts have to be secured to avoid risk of dropping. | FALSE
Start up instruction Protective earth measurements are described in start up instructions and done during system installation. | uring system installation | 19.38% | 24 | 990 | Start up instruction Protective earth measurements are described in start up instructions and done during system installation. | FALSE
Same as AT-S3-28789 - Missing cross-reference | reference | 20% | 325 | 809 | Same as AT-S3-28789 - Missing cross-reference | FALSE
Same as AT-S3-29472 - Missing cross-reference | reference | 20% | 325 | 482 | Same as AT-S3-29472 - Missing cross-reference | FALSE
Same as AT-S3-29480 - Missing cross-reference | reference | 20% | 325 | 1052 | Same as AT-S3-29480 - Missing cross-reference | FALSE
Same as AT-S3-28722 - Missing cross-reference | reference | 20% | 325 | 453 | Same as AT-S3-28722 - Missing cross-reference | FALSE
Same as AT-S3-28768 - Missing cross-reference | reference | 20% | 325 | 441 | Same as AT-S3-28768 - Missing cross-reference | FALSE
same as AT-S3-29198 - Missing cross-reference | reference | 20% | 325 | 440 | same as AT-S3-29198 - Missing cross-reference | FALSE
Same as AT-S3-29134 - Missing cross-reference | reference | 20% | 325 | 492 | Same as AT-S3-29134 - Missing cross-reference | FALSE
same as AT-S3-28444 - Missing cross-reference | reference | 20% | 325 | 1095 | same as AT-S3-28444 - Missing cross-reference | FALSE
Same as AT-S3-28732 - Missing cross-reference | reference | 20% | 325 | 422 | Same as AT-S3-28732 - Missing cross-reference | FALSE
CAUTION Caution: Position of emergency STOP buttons unknown. Possible consequences: Emergency STOP button will not be applied on time in case of an emergency! Remedy: Familiarize yourself with the position of emergency STOP buttons. You will find emergency STOP buttons in the following locations (see figure). | Emergency STOP button will not be applied on time in case of an | 19.44% | 29 | 899 | CAUTION Caution: Position of emergency STOP buttons unknown. Possible consequences: Emergency STOP button will not be applied on time in case of an emergency! Remedy: Familiarize yourself with the position of emergency STOP buttons. You will find emergency STOP buttons in the following locations (see figure). | TRUE
same as AT-S3-28768 - Missing cross-reference | reference | 20% | 325 | 437 | same as AT-S3-28768 - Missing cross-reference | FALSE
Same as AT-S3-29060 - Missing cross-reference | reference | 20% | 325 | 529 | Same as AT-S3-29060 - Missing cross-reference | FALSE
Same as AT-S3-29060 - Missing cross-reference | reference | 20% | 325 | 556 | Same as AT-S3-29060 - Missing cross-reference | FALSE
Same as AT-S3-28768 - Missing cross-reference | reference | 20% | 325 | 393 | Same as AT-S3-28768 - Missing cross-reference | FALSE
same as AT-S3-29060 - Missing cross-reference | reference | 20% | 325 | 394 | same as AT-S3-29060 - Missing cross-reference | FALSE
same as AT-S3-28437 - Missing cross-reference | reference | 20% | 325 | 1103 | same as AT-S3-28437 - Missing cross-reference | FALSE
Same as AT-S3-28789 - Missing cross-reference | reference | 20% | 325 | 421 | Same as AT-S3-28789 - Missing cross-reference | FALSE
same as AT-S3-28437 - Missing cross-reference | reference | 20% | 325 | 1100 | same as AT-S3-28437 - Missing cross-reference | FALSE
Same as AT-S3-28980 - Missing cross-reference | reference | 20% | 325 | 505 | Same as AT-S3-28980 - Missing cross-reference | FALSE
Regular check of the collimator plate adjustment per the Maintenance Plan by prescribed check. | of the collimator | 18.75% | 130 | 499 | Regular check of the collimator plate adjustment per the Maintenance Plan by prescribed check. | FALSE
CAUTION Cause: Increased temperature due to extended use Possible Consequences: Risk of burnings Remedy: The allowed temperature for brief contact is exceeded when the temperature switch of the single tank is triggered. | Increased temperature due to extended use | 18.3% | 328 | 724 | CAUTION Cause: Increased temperature due to extended use Possible Consequences: Risk of burnings Remedy: The allowed temperature for brief contact is exceeded when the temperature switch of the single tank is triggered. | TRUE
Attention note in service manual: Perform protective earth measurement after service work. | protective earth | 18.89% | 71 | 1191 | Attention note in service manual: Perform protective earth measurement after service work. | FALSE
same as AT-S3-29385 - Missing cross-reference | reference | 18.37% | 325 | 191 | same as AT-S3-29385 - Missing cross-reference | FALSE
Same as AT-S3-28598 - Missing cross-reference | reference | 19.15% | 325 | 1024 | Same as AT-S3-28598 - Missing cross-reference | FALSE
same as AT-S3-28443 - Missing cross-reference | reference | 20% | 325 | 1096 | same as AT-S3-28443 - Missing cross-reference | FALSE
Same as AT-S3-29060 - Missing cross-reference | reference | 20% | 325 | 540 | Same as AT-S3-29060 - Missing cross-reference | FALSE
Same as AT-S3-28397 - Missing cross-reference | reference | 20% | 325 | 559 | Same as AT-S3-28397 - Missing cross-reference | FALSE
Dielectric strength test shall be performed to guarantee proper insulation function. | ation function. | 17.86% | 189 | 985 | Dielectric strength test shall be performed to guarantee proper insulation function. | FALSE
Work step in maintenance documents: Perform protective earth measurement after maintenance work. | protective earth | 17.71% | 71 | 1192 | Work step in maintenance documents: Perform protective earth measurement after maintenance work. | FALSE
Work step in maintenance documents: operate leakage current measurement after maintenance work | fter maintenance | 18.09% | 48 | 1196 | Work step in maintenance documents: operate leakage current measurement after maintenance work | FALSE
The imaging system automatically performs a controlled shutdown, before the power of the Uninterrupted Power Supply runs out. | system automatically | 17.6% | 228 | 156 | The imaging system automatically performs a controlled shutdown, before the power of the Uninterrupted Power Supply runs out. | FALSE
same as AT-S3-29114 - Missing cross-reference | reference | 18.37% | 325 | 523 | same as AT-S3-29114 - Missing cross-reference | FALSE
A label shall be attached on Trolley housing and to unplug the power input connector for service. | rolley housing and | 18.45% | 299 | 248 | A label shall be attached on Trolley housing and to unplug the power input connector for service. | FALSE
Encapsulation and separate software tests of image flip/rotate. | ulation and | 18.46% | 320 | 678 | Encapsulation and separate software tests of image flip/rotate. | FALSE
No additional measures are required. The operators are trained on system behavior('advice to Fluoro exposure with > 2fps before acquiring a single image in order to set the exposure parameters and to achieve good image quality). | e in order to set the exposure parameters | 18.42% | 129 | 625 | No additional measures are required. The operators are trained on system behavior('advice to Fluoro exposure with > 2fps before acquiring a single image in order to set the exposure parameters and to achieve good image quality). | FALSE
Visibility of the top of K-Wire in an appropriate way (e.g. only transparent overlay). | sibility of the | 18.6% | 15 | 657 | Visibility of the top of K-Wire in an appropriate way (e.g. only transparent overlay). | FALSE
Same as AT-S3-28641 - Missing cross-reference | reference | 19.15% | 325 | 981 | Same as AT-S3-28641 - Missing cross-reference | FALSE
Protective cover should be added into the screw arbor of the transformer | of the transf | 19.44% | 220 | 971 | Protective cover should be added into the screw arbor of the transformer | FALSE
When storage capacity is exhausted, the image system signals 'disk full, radiation still possible' | storage capacity | 18.37% | 276 | 654 | When storage capacity is exhausted, the image system signals 'disk full, radiation still possible' | FALSE
see AT-S3-28515 - Missing cross-reference | reference | 18.75% | 325 | 1106 | see AT-S3-28515 - Missing cross-reference | FALSE
CAUTION Cause: Incorrect grid status: inserted, removed. Possible consequences: X-ray exposure without clinical benefit. Remedy: - Select the application according to your clinical procedure. - Make sure that the current grid status meets your image requirements. The imaging system indicates whether a grid has been inserted or removed. - Insert or remove the grid manually. | Make sure that the current grid status meets your image requirements. | 18.37% | 109 | 512 | CAUTION Cause: Incorrect grid status: inserted, removed. Possible consequences: X-ray exposure without clinical benefit. Remedy: - Select the application according to your clinical procedure. - Make sure that the current grid status meets your image requirements. The imaging system indicates whether a grid has been inserted or removed. - Insert or remove the grid manually. | TRUE
Same as AT-S3-28411 - Missing cross-reference | reference | 19.15% | 325 | 967 | Same as AT-S3-28411 - Missing cross-reference | FALSE
The design of the external accessible connectors must consider misplugging of cables. Compare also chapter 15.4.1 of IEC60601-1 | accessible connectors | 17.05% | 272 | 1084 | The design of the external accessible connectors must consider misplugging of cables. Compare also chapter 15.4.1 of IEC60601-1 | FALSE
Warning labels 'risk of pinch - hand' shall be placed on the angulation cover (upper and lower side), for orbital movement. | for orbital movement | 16.67% | 81 | 916 | Warning labels 'risk of pinch - hand' shall be placed on the angulation cover (upper and lower side), for orbital movement. | FALSE
Monitor arm shall be designed in a way so that an overload can be carried in compliance with IEC 60601-1, chap. 9.8., table 21 | ance with IEC 60601-1 | 16.41% | 300 | 927 | Monitor arm shall be designed in a way so that an overload can be carried in compliance with IEC 60601-1, chap. 9.8., table 21 | FALSE
Attention note in the relevant service manual: Operate leakage current measurement after service. | the relevant se | 16.49% | 185 | 1195 | Attention note in the relevant service manual: Operate leakage current measurement after service. | FALSE
Install additional instruction plate. __TBI__Please read user manual__TBI__ | additional i | 17.33% | 221 | 1017 | Install additional instruction plate. __TBI__Please read user manual__TBI__ | FALSE
Factory documents: A warning label 'Hot surface' shall be put at the tube housing. | warning label | 18.07% | 292 | 962 | Factory documents: A warning label 'Hot surface' shall be put at the tube housing. | FALSE
No additional measures are required. The operators are trained on system behavior and correct operation / workflow. | additional measures | 18.26% | 49 | 623 | No additional measures are required. The operators are trained on system behavior and correct operation / workflow. | FALSE
CAUTION Cause: Abrasion of wear. Possible consequences: Injury of patient, operating personnel and others. Permanent damage to the system. Remedy: Consider the maintenance guidelines to maintain safety and proper function of the system. | the maintenance guidelines to maintain | 16.6% | 23 | 1001 | CAUTION Cause: Abrasion of wear. Possible consequences: Injury of patient, operating personnel and others. Permanent damage to the system. Remedy: Consider the maintenance guidelines to maintain safety and proper function of the system. | TRUE
If the temperature within the mono block exceeds 50°C a warning signal will be initiated to the user. | f the temperature | 17.48% | 140 | 965 | If the temperature within the mono block exceeds 50°C a warning signal will be initiated to the user. | FALSE
Imaging system shall ensure that all previously acquired data remain uncorrupted. | maging system s | 18.07% | 90 | 1067 | Imaging system shall ensure that all previously acquired data remain uncorrupted. | FALSE
CAUTION Cause: Tripping over cables Possible consequences: Risk to fall Remedy: Installing cables, make sure they lay neatly on the ground and that there are no loops. | ing cables, make sure they | 15.88% | 69 | 1090 | CAUTION Cause: Tripping over cables Possible consequences: Risk to fall Remedy: Installing cables, make sure they lay neatly on the ground and that there are no loops. | TRUE
CAUTION Cause: The system is not designed for operation in explosion-endangered areas. It does not fulfill the requirements for AP/APG classification. Possible consequences: Danger of explosion Remedy: Do not use the system where explosive atmospheres can arise. | The system is not designed for operation in | 16.48% | 329 | 1213 | CAUTION Cause: The system is not designed for operation in explosion-endangered areas. It does not fulfill the requirements for AP/APG classification. Possible consequences: Danger of explosion Remedy: Do not use the system where explosive atmospheres can arise. | TRUE
The wheels must be designed in a way that an overload can be carried in compliance with IEC 60601-1 chap. 9.8, table 21 | ance with IEC 60601-1 | 17.36% | 300 | 1250 | The wheels must be designed in a way that an overload can be carried in compliance with IEC 60601-1 chap. 9.8, table 21 | FALSE
The emergency stop circuit shall be checked during start-up. In case the test is not passed, motorized movements shall remain blocked. | motorized movements | 15.67% | 74 | 802 | The emergency stop circuit shall be checked during start-up. In case the test is not passed, motorized movements shall remain blocked. | FALSE
The welded joints must be designed in a way so that an overload can be carried in compliance with IEC 60601-1 chap. 9.8., table 21 | ance with IEC 60601-1 | 15.91% | 300 | 1253 | The welded joints must be designed in a way so that an overload can be carried in compliance with IEC 60601-1 chap. 9.8., table 21 | FALSE
CAUTION Cause: Nonconforming wrapping of the C-arm with the sterile cover. Possible consequences: Risk of infection! Remedy: Observe the instructions for handling of the sterile cover. | the instructions for handling | 16.58% | 307 | 731 | CAUTION Cause: Nonconforming wrapping of the C-arm with the sterile cover. Possible consequences: Risk of infection! Remedy: Observe the instructions for handling of the sterile cover. | TRUE
Radiation is switched off after a maximum mAs(Radigraph), Switch off radiation after max time (fluoro). | is switched off a | 17.48% | 111 | 531 | Radiation is switched off after a maximum mAs(Radigraph), Switch off radiation after max time (fluoro). | FALSE
CAUTION Cause: Wrong or extreme image processing parameters. Possible consequences: Risk of limited image quality in the resulting X-ray images. Risk of incorrect decision during intervention or artifacts. Remedy: - Apply image processing parameters carefully, such as edge enhancement and windowing, and always check them carefully. - Always check the image quality before reaching an interventional decision. | Risk of limited image quality in the resulting X-ray images. Risk of | 16.55% | 108 | 94 | CAUTION Cause: Wrong or extreme image processing parameters. Possible consequences: Risk of limited image quality in the resulting X-ray images. Risk of incorrect decision during intervention or artifacts. Remedy: - Apply image processing parameters carefully, such as edge enhancement and windowing, and always check them carefully. - Always check the image quality before reaching an interventional decision. | TRUE
Guarantee the Ethernet cable (from MCU to generator) is strang (not easy to shake) and signal stable. | Ethernet cable | 15.84% | 301 | 168 | Guarantee the Ethernet cable (from MCU to generator) is strang (not easy to shake) and signal stable. | FALSE
Use battery cases or housings that prevent an accidental short-circuit | e battery ca | 16.67% | 302 | 198 | Use battery cases or housings that prevent an accidental short-circuit | FALSE
CAUTION Cause: Detector temperature out of range Possible consequences: Risk of limited image quality in resulting X-ray images without value for interventional decision. Remedy: Check the temperature icon on the display regularly. | heck the temperature icon on the display | 16.95% | 108 | 100 | CAUTION Cause: Detector temperature out of range Possible consequences: Risk of limited image quality in resulting X-ray images without value for interventional decision. Remedy: Check the temperature icon on the display regularly. | TRUE
After the activation of the mechanical safety feature (trap nut), the operation of the drive must not be possible anymore. | f the mechanical safe | 16.94% | 44 | 776 | After the activation of the mechanical safety feature (trap nut), the operation of the drive must not be possible anymore. | FALSE
CAUTION Cause: Image orientation has been reset Possible consequences: Risk of wrong interventional decision Remedy: Duty of care for operator, that the image orientation will be rest after an application change | Risk of wrong interventional decision | 17.29% | 156 | 666 | CAUTION Cause: Image orientation has been reset Possible consequences: Risk of wrong interventional decision Remedy: Duty of care for operator, that the image orientation will be rest after an application change | TRUE
The temperature of the monoblock shall be monitored. X-ray shall be inhibited when the maximum operating temperature is reached. see also AT-S3-28911 - Missing cross-reference. | temperature of the monobloc | 15.82% | 32 | 725 | The temperature of the monoblock shall be monitored. X-ray shall be inhibited when the maximum operating temperature is reached. see also AT-S3-28911 - Missing cross-reference. | FALSE
Make sure that the date and number formats as well as distance units (inch, cm/mm) are correctly configured for the customer. | Make sure that the d | 16% | 110 | 406 | Make sure that the date and number formats as well as distance units (inch, cm/mm) are correctly configured for the customer. | FALSE
Generator shall send off an signal to indicate whether X ray radiation is on or not. | radiation is | 16.67% | 140 | 635 | Generator shall send off an signal to indicate whether X ray radiation is on or not. | FALSE
Supervision of power in PowerUnit: The PowerUnit shall provide a protection against excessive current and voltage. | protection against | 17.09% | 19 | 701 | Supervision of power in PowerUnit: The PowerUnit shall provide a protection against excessive current and voltage. | FALSE
Perform endurance test for SG cable and C arm cable. Perform drop test for SG cable, according to IEC60601-1. | according to IEC | 15.45% | 275 | 1118 | Perform endurance test for SG cable and C arm cable. Perform drop test for SG cable, according to IEC60601-1. | FALSE
Even battery is exhaused during standby mode, system can start up normally after power line connected and system switched on. | power line connect | 15.2% | 71 | 387 | Even battery is exhaused during standby mode, system can start up normally after power line connected and system switched on. | FALSE
An instrusion detection solution is provided by the system, for e.g., based on solidecore/whitelist. | provided by the | 16% | 298 | 376 | An instrusion detection solution is provided by the system, for e.g., based on solidecore/whitelist. | FALSE
Consistency check in imaging pipeline and notification of host in case of error. | case of error. | 16.87% | 218 | 605 | Consistency check in imaging pipeline and notification of host in case of error. | FALSE
Test of UI (text strings and symbols) in each language provided. | age provide | 17.19% | 18 | 417 | Test of UI (text strings and symbols) in each language provided. | FALSE
Warning Cause: Intrusion of liquids Possible Consequences: Damage of FD Remedy: Ensure that FD is covered by drape at any time during intervention. | e during intervention. | 14.67% | 106 | 688 | Warning Cause: Intrusion of liquids Possible Consequences: Damage of FD Remedy: Ensure that FD is covered by drape at any time during intervention. | TRUE
The system is designed in a way that prevents tipping over on an incline at an angle of 10°. Remark: Compliance with IEC 60601-1 tilt test | ance with IEC 60601-1 | 15% | 300 | 1164 | The system is designed in a way that prevents tipping over on an incline at an angle of 10°. Remark: Compliance with IEC 60601-1 tilt test | FALSE
CAUTION Cause: Exposure does not stop due to a technical defect or unintentional exposure. Possible consequences: Risk of unnecessary radiation exposure Remedy: In case of unwanted radiation, press the nearest emergency STOP button, for the case it is not pressed already. | of unwanted radiation, press the nearest | 15.27% | 36 | 1079 | CAUTION Cause: Exposure does not stop due to a technical defect or unintentional exposure. Possible consequences: Risk of unnecessary radiation exposure Remedy: In case of unwanted radiation, press the nearest emergency STOP button, for the case it is not pressed already. | TRUE
Imaging system assigns acquired 2D images to the patient who was registered at the time the images have been acquired. | the time the imag | 15.25% | 230 | 463 | Imaging system assigns acquired 2D images to the patient who was registered at the time the images have been acquired. | FALSE
CAUTION Cause: Usage of improper cleaning agents. Possible consequences: Risk of contamination! Remedy: Use only the recommended agents for cleaning and disinfection. | cleaning and disinfection | 15.38% | 254 | 449 | CAUTION Cause: Usage of improper cleaning agents. Possible consequences: Risk of contamination! Remedy: Use only the recommended agents for cleaning and disinfection. | TRUE
A pressure compensation membrane for oil expansion has to be provided for on the monoblock. | ion has to be | 15.38% | 239 | 721 | A pressure compensation membrane for oil expansion has to be provided for on the monoblock. | FALSE
CAUTION Cause: Reduced image quality due to filtering Possible consequences: Risk of artifacts e.g. metal implants can appear to be loosening phenomena. Filtering may cause loss of image quality. Remedy: Choose the Gain value, LUT and low pass filter carefully. | metal implants can appear to be loosening | 15.73% | 144 | 519 | CAUTION Cause: Reduced image quality due to filtering Possible consequences: Risk of artifacts e.g. metal implants can appear to be loosening phenomena. Filtering may cause loss of image quality. Remedy: Choose the Gain value, LUT and low pass filter carefully. | TRUE
Specify smooth surface finishes (referring to IEC60601-1 chapter 9.3.) | IEC60601-1 | 15.49% | 271 | 1176 | Specify smooth surface finishes (referring to IEC60601-1 chapter 9.3.) | FALSE
The imaging system hardware shall be tested at startup, including the connection to the detector. In case an error is detected, the imaging system shall not report ready state. | e imaging system hardware | 13.97% | 325 | 603 | The imaging system hardware shall be tested at startup, including the connection to the detector. In case an error is detected, the imaging system shall not report ready state. | FALSE
CAUTION Cause: Image orientation mixed up Possible consequences: Risk of wrong intraoperative decision Remedy: Duty of care for operator; recommendation to label each side. | Duty of care for operator | 14.86% | 156 | 683 | CAUTION Cause: Image orientation mixed up Possible consequences: Risk of wrong intraoperative decision Remedy: Duty of care for operator; recommendation to label each side. | TRUE
CAUTION Cause: Movement without intentional activation of control elements. Possible consequences: Risk of collision, risk of injury to patient or operator, risk of damage to unit parts. Remedy: If movement does not stop press the nearest emergency STOP button. | Risk of collision, risk of injury to | 14.02% | 74 | 843 | CAUTION Cause: Movement without intentional activation of control elements. Possible consequences: Risk of collision, risk of injury to patient or operator, risk of damage to unit parts. Remedy: If movement does not stop press the nearest emergency STOP button. | TRUE
The imaging system hardware shall be tested at startup, including the connection to the detector. In case an error is detected, the imaging system shall not report ready state. | e imaging system hardware | 14.2% | 325 | 627 | The imaging system hardware shall be tested at startup, including the connection to the detector. In case an error is detected, the imaging system shall not report ready state. | FALSE
Usage of save hardware modules and signal transmission line (e.g redundant hardware contacts and/or CRC for transmission line) | for transmission | 14.06% | 148 | 348 | Usage of save hardware modules and signal transmission line (e.g redundant hardware contacts and/or CRC for transmission line) | FALSE
CAUTION Cause: Weak batteries of wireless foot switch. Risk: Discontinuation of Interventional treatment. Remedy: Make sure the batteries have enough electricity prior to starting an examination. | ake sure the batteries have | 14% | 305 | 340 | CAUTION Cause: Weak batteries of wireless foot switch. Risk: Discontinuation of Interventional treatment. Remedy: Make sure the batteries have enough electricity prior to starting an examination. | TRUE
Usage of a battery with the possibility to work with wireless footswitch. The user shall be informed via LEDs about low capacity to change the battery. | wireless footswitch. | 13.91% | 325 | 339 | Usage of a battery with the possibility to work with wireless footswitch. The user shall be informed via LEDs about low capacity to change the battery. | FALSE
Develop design without sharp edges. (referring to IEC60601-1 chapter 9.3.) | IEC60601-1 | 14.67% | 271 | 698 | Develop design without sharp edges. (referring to IEC60601-1 chapter 9.3.) | FALSE
The wireless and the wired footswitch shall have protection class IPx8 according to EN 60529 with closed battery compartment. | battery compartment | 15.62% | 305 | 618 | The wireless and the wired footswitch shall have protection class IPx8 according to EN 60529 with closed battery compartment. | FALSE
The monitor support must be designed in a way that an overload can be carried. Remark: Compliance with IEC 60601-1 chapter 9.8 Table 21 | ance with IEC 60601-1 | 15.11% | 300 | 1200 | The monitor support must be designed in a way that an overload can be carried. Remark: Compliance with IEC 60601-1 chapter 9.8 Table 21 | FALSE
The moving positioning units must be designed so that an overload can be carried in compliance with IEC 60601-1, Chap. 9.8, Table 21. | ance with IEC 60601-1 | 15.56% | 300 | 781 | The moving positioning units must be designed so that an overload can be carried in compliance with IEC 60601-1, Chap. 9.8, Table 21. | FALSE
Put the one 'No sitting' warning label on detector top surface. Put the two 'No stepping on surface' warning labels on both sides of the tube. | on both sides of the | 15.17% | 94 | 735 | Put the one 'No sitting' warning label on detector top surface. Put the two 'No stepping on surface' warning labels on both sides of the tube. | FALSE
The brakes must be designed with sufficient braking power. Remark: This is in compliance with IEC 60601-1; clause 9.4.3 (Instability). | ance with IEC 60601-1 | 15.22% | 300 | 1245 | The brakes must be designed with sufficient braking power. Remark: This is in compliance with IEC 60601-1; clause 9.4.3 (Instability). | FALSE
CAUTION Cause: Patient ID already exists. Possible Consequences: Misinterpretation of data of the examined patient if the selection is based solely on the patient ID. Remedy: -Always use four-way identification [patient name, sex, date of birth, patient ID].- Avoid mixing up patient data by using four-way identification. Selecting patient data based on the patient ID alone is insufficient. Ensure the proper display of parameters in the Worklist and in the Patient list. | Avoid mixing up patient data by using four-way identification. Selecting | 15.3% | 115 | 584 | CAUTION Cause: Patient ID already exists. Possible Consequences: Misinterpretation of data of the examined patient if the selection is based solely on the patient ID. Remedy: -Always use four-way identification [patient name, sex, date of birth, patient ID].- Avoid mixing up patient data by using four-way identification. Selecting patient data based on the patient ID alone is insufficient. Ensure the proper display of parameters in the Worklist and in the Patient list. | TRUE
Label with book reader symbol and symbol to drive the vertical column completely downwards. | the vertical | 15.38% | 80 | 761 | Label with book reader symbol and symbol to drive the vertical column completely downwards. | FALSE
CAUTION: Cause: In extreme operating mode (continuous operation), the temperature of the X-ray tube assembly can heat up. Possible consequences: Danger of burns. Remedy: Avoid contact with the tube housing | (continuous operation), the | 13.94% | 32 | 961 | CAUTION: Cause: In extreme operating mode (continuous operation), the temperature of the X-ray tube assembly can heat up. Possible consequences: Danger of burns. Remedy: Avoid contact with the tube housing | TRUE
In addition to the SW control the activation of any motorized movements has to be secured by a separate independent HW circuit e.g. current loop. | motorized movements | 14.29% | 74 | 738 | In addition to the SW control the activation of any motorized movements has to be secured by a separate independent HW circuit e.g. current loop. | FALSE
Attach label with the inscription: Warning: Voltage present. Any exchange by staff not sufficiently trained can be dangerous. | not sufficiently | 14.29% | 10 | 987 | Attach label with the inscription: Warning: Voltage present. Any exchange by staff not sufficiently trained can be dangerous. | FALSE
The trolley must be designed in a way that prevents tipping over when placed in any transport position on an inclined plane at an angle of 10 °. | transport position | 13.89% | 92 | 1274 | The trolley must be designed in a way that prevents tipping over when placed in any transport position on an inclined plane at an angle of 10 °. | FALSE
After reaching the safety distance an acoustic alarm signal is triggered, while moving C-arm downwards in direction of crushing hazard. | the safety distance | 14.81% | 326 | 798 | After reaching the safety distance an acoustic alarm signal is triggered, while moving C-arm downwards in direction of crushing hazard. | FALSE
Self-test of emergency stop circuit at startup. If fail, then block movement only. radiation is still allowed. | emergency stop | 14.29% | 98 | 741 | Self-test of emergency stop circuit at startup. If fail, then block movement only. radiation is still allowed. | FALSE
A labeling, safety sign ISO 7010-W001, has to be placed next to the mechanical safety equipment or in a way so the connection to it is obvious. | the mechanical safe | 13.99% | 44 | 774 | A labeling, safety sign ISO 7010-W001, has to be placed next to the mechanical safety equipment or in a way so the connection to it is obvious. | FALSE
If the temperature exceeds 70°C the ongoing x-ay radiation will be stopped and a new radiation request will be blocked. | radiation will be | 14.88% | 150 | 964 | If the temperature exceeds 70°C the ongoing x-ay radiation will be stopped and a new radiation request will be blocked. | FALSE
The system must be designed in a way that prevents tipping over on a incline at an angle of 5°. Remark: Compliance with IEC 60601-1 tilt test | ance with IEC 60601-1 | 14.69% | 300 | 1168 | The system must be designed in a way that prevents tipping over on a incline at an angle of 5°. Remark: Compliance with IEC 60601-1 tilt test | FALSE
CAUTION Cause: In extreme operating mode (continuous operation), the temperature of the X-ray tube assembly can heat up. Possible consequences: Danger of burns. Remedy: Avoid contact with the tube housing | (continuous operation), the | 14.01% | 32 | 960 | CAUTION Cause: In extreme operating mode (continuous operation), the temperature of the X-ray tube assembly can heat up. Possible consequences: Danger of burns. Remedy: Avoid contact with the tube housing | FALSE
During operating of the system there will be a self test of the temperature sensing circuit. In the case of a defect x-ray radiation will be prevented. | ting of the system | 12.42% | 66 | 963 | During operating of the system there will be a self test of the temperature sensing circuit. In the case of a defect x-ray radiation will be prevented. | FALSE
Confirmation by user when deleting studies where nothing has been archived or printed. Message: 'Do you want unarchived studies to be deleted?' | has been archived | 13.29% | 169 | 864 | Confirmation by user when deleting studies where nothing has been archived or printed. Message: 'Do you want unarchived studies to be deleted?' | FALSE
Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a link speed detection. | able data transfer | 13.77% | 219 | 572 | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a link speed detection. | FALSE
No hazard on system level Note: A delay time of max. 200 msec is tolerable to ensure sufficient reaction time for the operator during intervention procedures. | during intervention | 13.21% | 183 | 1045 | No hazard on system level Note: A delay time of max. 200 msec is tolerable to ensure sufficient reaction time for the operator during intervention procedures. | FALSE
see Risk Analysis - Identified Risk Control Measures - AT-S3-27124 - User Manual - Calibration procedure Note: VA21 is only suitable for FD version. | alibration procedure | 13.42% | 190 | 430 | see Risk Analysis - Identified Risk Control Measures - AT-S3-27124 - User Manual - Calibration procedure Note: VA21 is only suitable for FD version. | FALSE
Pressure switch in tube is monitored during exposure. If excess pressure is detected, the switch is opened and the radiation is switched off. | and the radiation i | 14.18% | 26 | 955 | Pressure switch in tube is monitored during exposure. If excess pressure is detected, the switch is opened and the radiation is switched off. | FALSE
Pressure switch in tube is monitored during exposure. If excess pressure is detected, the switch is opened and the radiation is switched off. | and the radiation i | 14.18% | 26 | 719 | Pressure switch in tube is monitored during exposure. If excess pressure is detected, the switch is opened and the radiation is switched off. | FALSE
Virus scanner is configured at installation time in the following way: - scan only incoming files on certain directories - no scheduled scan over all files (only real time scan) | in the following way: | 12.15% | 120 | 380 | Virus scanner is configured at installation time in the following way: - scan only incoming files on certain directories - no scheduled scan over all files (only real time scan) | FALSE
see AT-S3-28551 - Missing cross-reference and AT-S3-28552 - Missing cross-reference | reference | 12.05% | 310 | 472 | see AT-S3-28551 - Missing cross-reference and AT-S3-28552 - Missing cross-reference | FALSE
CAUTION Cause: The Wireless Footswitch may be interfered by other equipment, including portable and fixed RF communication equipment, even if such equipment meets the applicable emissions requirements. Risk: Fluoro or Acquisition may sporadically be interrupted. Remedy: The operator must assure that other wireless device in the 2,4 GHz ISM band should not be operated in the vicinity of appr. 5m around the S3P System. Please observe and verify normal operation of the Wireless Footswitch prior to using it. | equipment, including portable and fixed RF communication equipment, | 13.04% | 303 | 342 | CAUTION Cause: The Wireless Footswitch may be interfered by other equipment, including portable and fixed RF communication equipment, even if such equipment meets the applicable emissions requirements. Risk: Fluoro or Acquisition may sporadically be interrupted. Remedy: The operator must assure that other wireless device in the 2,4 GHz ISM band should not be operated in the vicinity of appr. 5m around the S3P System. Please observe and verify normal operation of the Wireless Footswitch prior to using it. | TRUE
WARNING Cause: Manual or motorized movements Possible consequences: Risks of injury by crushing or collision for persons and objects in the vicinity of moving parts. Remedy: Please ensure during either manual or motorized movements of the system, that collisions with persons or other objects, which are in system's path, are prevented at any time. Ensure that everybody is outside the hazardous zone. | collision for persons and objects in the vicinity of | 13.05% | 74 | 918 | WARNING Cause: Manual or motorized movements Possible consequences: Risks of injury by crushing or collision for persons and objects in the vicinity of moving parts. Remedy: Please ensure during either manual or motorized movements of the system, that collisions with persons or other objects, which are in system's path, are prevented at any time. Ensure that everybody is outside the hazardous zone. | FALSE
CAUTION Cause: Wrong exposure parameters Possible consequences: Risk of unnecessary X-ray exposure due to wrong exposure parameters. Remedy: Take care of the correct exposure parameters. Pay attention to the exposure index (EXI). | correct exposure parameters. | 11.97% | 329 | 524 | CAUTION Cause: Wrong exposure parameters Possible consequences: Risk of unnecessary X-ray exposure due to wrong exposure parameters. Remedy: Take care of the correct exposure parameters. Pay attention to the exposure index (EXI). | TRUE
CAUTION Cause: Accidentally step on foot switch Possible consequences: Risk of radiation Remedy: Switch-off system before starting cleaning procedure | cleaning procedure | 12.34% | 255 | 473 | CAUTION Cause: Accidentally step on foot switch Possible consequences: Risk of radiation Remedy: Switch-off system before starting cleaning procedure | TRUE
Hand and foot switches have mechanical or electrical measures to help prevent accidental activation. | e mechanical | 13% | 44 | 636 | Hand and foot switches have mechanical or electrical measures to help prevent accidental activation. | FALSE
Define a separate testcase and perform test of UI (text strings and symbols) for each foreign user language. | e and perform | 12.04% | 24 | 415 | Define a separate testcase and perform test of UI (text strings and symbols) for each foreign user language. | FALSE
CAUTION Cause: Wrong exposure parameters Possible consequences: Risk of unnecessary X-ray exposure due to wrong exposure parameters. Remedy: Take care of the correct exposure parameters. Pay attention to the exposure index (EXI). | correct exposure parameters. | 12.07% | 329 | 520 | CAUTION Cause: Wrong exposure parameters Possible consequences: Risk of unnecessary X-ray exposure due to wrong exposure parameters. Remedy: Take care of the correct exposure parameters. Pay attention to the exposure index (EXI). | FALSE
CAUTION Cause: Use of harsh cleaning agents, liquids or sprays. Possible consequences: Risk of electrical hazard or damage to the system Remedy: Use only substances for cleaning and disinfection, which are recommended. Do not let cleaning liquids seep into the openings of the system (e.g. air openings, gaps between covers). Observe the cleaning and disinfection instructions. | Risk of electrical hazard or damage to the system | 12.89% | 256 | 476 | CAUTION Cause: Use of harsh cleaning agents, liquids or sprays. Possible consequences: Risk of electrical hazard or damage to the system Remedy: Use only substances for cleaning and disinfection, which are recommended. Do not let cleaning liquids seep into the openings of the system (e.g. air openings, gaps between covers). Observe the cleaning and disinfection instructions. | TRUE
Access to BIOS needs to be password-protected. Note: Subsystem APPhost is only suitable for VA21 onword. | suitable for | 13.08% | 310 | 372 | Access to BIOS needs to be password-protected. Note: Subsystem APPhost is only suitable for VA21 onword. | FALSE
CAUTION Cause: Restart missing after Service session Possible consequences: Changes in the Service environment may not be effective without a restart of the system. Remedy: Restart the system after a Service session. | Restart the system after a | 12.67% | 248 | 83 | CAUTION Cause: Restart missing after Service session Possible consequences: Changes in the Service environment may not be effective without a restart of the system. Remedy: Restart the system after a Service session. | TRUE
CAUTION Cause: Restart missing after Service session Possible consequences: Changes in the Service environment may not be effective without a restart of the system. Remedy: Restart the system after a Service session. | Restart the system after a | 12.79% | 248 | 46 | CAUTION Cause: Restart missing after Service session Possible consequences: Changes in the Service environment may not be effective without a restart of the system. Remedy: Restart the system after a Service session. | FALSE
CAUTION Cause: Brake failure Possible consequences: Risk of crushing Remedy: Before starting the treatment, perform daily function and safety checks. | and safety checks. | 12.5% | 242 | 1247 | CAUTION Cause: Brake failure Possible consequences: Risk of crushing Remedy: Before starting the treatment, perform daily function and safety checks. | TRUE
In order to avoid hazards from fire or smoke by overheating or flammable materials the system has to be developed in accordance with the IEC 60601-1 standard. Note: Subsystem AppHost only suitable for system with it. | in accordance with the IEC | 12.22% | 234 | 1219 | In order to avoid hazards from fire or smoke by overheating or flammable materials the system has to be developed in accordance with the IEC 60601-1 standard. Note: Subsystem AppHost only suitable for system with it. | FALSE
Cause: Defective Imaging system Hardware Possible consequences: Risk of X-ray radiation without a stored image Remedy: In case of image failures stop exposures and switch off the system | Risk of X-ray radiation | 12.57% | 139 | 97 | Cause: Defective Imaging system Hardware Possible consequences: Risk of X-ray radiation without a stored image Remedy: In case of image failures stop exposures and switch off the system | FALSE
Warning Cause: Intrusion of instruments into gaps between control panel and control panel housing Possible Consequences: Short circuit, electric shock, and/or fire Remedy: Do not put neither instruments nor conductive material nor liquids into openings/gaps of control panel and control panel housing. | hort circuit, electric shock, and/or fire | 13.4% | 107 | 25 | Warning Cause: Intrusion of instruments into gaps between control panel and control panel housing Possible Consequences: Short circuit, electric shock, and/or fire Remedy: Do not put neither instruments nor conductive material nor liquids into openings/gaps of control panel and control panel housing. | TRUE
In order to prevent electrostatic charging (e. g. when the system is moved over the floor), the system shall provide a conductive connection to the ground, e.g. conductance band, conductive wheels. | a conductive connection | 12.56% | 33 | 1187 | In order to prevent electrostatic charging (e. g. when the system is moved over the floor), the system shall provide a conductive connection to the ground, e.g. conductance band, conductive wheels. | FALSE
Service Safety Instruction CAUTION Cause: Usage of wrong fuses Possible Consequences: Risk of fire Remedy: Ensure the correct usage of fuses taking into account the label next to the fuses. | taking into account the | 12.76% | 209 | 119 | Service Safety Instruction CAUTION Cause: Usage of wrong fuses Possible Consequences: Risk of fire Remedy: Ensure the correct usage of fuses taking into account the label next to the fuses. | FALSE
Monobloc needs to be covered with housing to prohibit direct contact of user/patients to parts of monoblock. | covered with | 12.96% | 306 | 727 | Monobloc needs to be covered with housing to prohibit direct contact of user/patients to parts of monoblock. | FALSE
Same as Risk Analysis - Identified Risk Control Measures - AT-S3-27295 - User Manual-EmergencyStop | EmergencyStop | 13.27% | 329 | 840 | Same as Risk Analysis - Identified Risk Control Measures - AT-S3-27295 - User Manual-EmergencyStop | FALSE
same as Risk Analysis - Identified Risk Control Measures - AT-S3-27848 - OperatorManual-Virus infection | OperatorManual | 13.21% | 329 | 378 | same as Risk Analysis - Identified Risk Control Measures - AT-S3-27848 - OperatorManual-Virus infection | FALSE
The user manual should have the guidline to consult the user to exchange the batteries. Note: This is suitable for the system with wireless footswitch. | wireless footswitch. | 13.46% | 325 | 200 | The user manual should have the guidline to consult the user to exchange the batteries. Note: This is suitable for the system with wireless footswitch. | FALSE
same as Risk Analysis - Identified Risk Control Measures - AT-S3-26994 - Labeling according to XPQR 4.8/01 | according to | 13.21% | 320 | 205 | same as Risk Analysis - Identified Risk Control Measures - AT-S3-26994 - Labeling according to XPQR 4.8/01 | FALSE
In operation modes for which saving are configured, a warning message is displayed that saving is not possible. Radiation with LIH-function is possible nonetheless. | message is displayed | 13.41% | 242 | 879 | In operation modes for which saving are configured, a warning message is displayed that saving is not possible. Radiation with LIH-function is possible nonetheless. | FALSE
Warning labeling Warning labeling at the power inlet of the image system. Unplug the power input connector for service. | he image system | 12.2% | 11 | 249 | Warning labeling Warning labeling at the power inlet of the image system. Unplug the power input connector for service. | FALSE
Use only certified power cable (e.g. UL Style 20235) to reduce risk of damage of the cable during normal system use | risk of damage | 13.04% | 74 | 1119 | Use only certified power cable (e.g. UL Style 20235) to reduce risk of damage of the cable during normal system use | FALSE
CAUTION Cause: Radiation indicator fails Possible consequences: Risk of unintentional radiation Remedy: Please call service provider if a radiation indicator has failed. | adiation indicator fails | 13.79% | 36 | 1075 | CAUTION Cause: Radiation indicator fails Possible consequences: Risk of unintentional radiation Remedy: Please call service provider if a radiation indicator has failed. | TRUE
If Emergency-STOP pushbutton is pressed all motorized movements are stopped. The Emergency-STOP button is installed within easy reach at all control locations. | motorized movements | 13.21% | 74 | 844 | If Emergency-STOP pushbutton is pressed all motorized movements are stopped. The Emergency-STOP button is installed within easy reach at all control locations. | FALSE
The storage rate for radiography shall be set to 100% and shall not be configurable by service or operator. | he storage rate | 13.64% | 179 | 592 | The storage rate for radiography shall be set to 100% and shall not be configurable by service or operator. | FALSE
During system startup, PM in MCU will trigger selftest of generator protective path. Only after test pass, intended radiation is allowed. | uring system start | 12.95% | 72 | 1078 | During system startup, PM in MCU will trigger selftest of generator protective path. Only after test pass, intended radiation is allowed. | FALSE
CAUTION Cause: Due to the complexity of the system or line voltage below the specified threshold, the loss of X-ray imaging or other system functions during an exam cannot be completely excluded. Possible consequences: Delayed or no imaging, possibly suboptimal treatment. Remedy: Therefore consider the need to establish emergency procedures in such cases. | , the loss of X-ray imaging or other system | 11.94% | 27 | 1228 | CAUTION Cause: Due to the complexity of the system or line voltage below the specified threshold, the loss of X-ray imaging or other system functions during an exam cannot be completely excluded. Possible consequences: Delayed or no imaging, possibly suboptimal treatment. Remedy: Therefore consider the need to establish emergency procedures in such cases. | FALSE
CATION Cause: Risk that patients are mixed-up. Possible consequences: Wrong patient is treated. Remedy: The user has to assure that the correct patient is treated before he is registered at the system. | the correct patient is | 11.65% | 114 | 79 | CATION Cause: Risk that patients are mixed-up. Possible consequences: Wrong patient is treated. Remedy: The user has to assure that the correct patient is treated before he is registered at the system. | TRUE
Use of unauthorized software or manipulations and alterations of the software can cause the product to function incorrectly and cause injury to the patient and damage to the product. Use software with correct version approved by Siemens. | injury to the patient and | 11.3% | 33 | 275 | Use of unauthorized software or manipulations and alterations of the software can cause the product to function incorrectly and cause injury to the patient and damage to the product. Use software with correct version approved by Siemens. | FALSE
Caution in the user manual Cause: Defective FLC-Hardware Possible consequences: Risk of X-ray radiation without a stored image Remedy: In case of image failures stop exposures and switch off the x ray unit | Risk of X-ray radiation | 11.43% | 139 | 60 | Caution in the user manual Cause: Defective FLC-Hardware Possible consequences: Risk of X-ray radiation without a stored image Remedy: In case of image failures stop exposures and switch off the x ray unit | FALSE
CAUTION Cause: Risk of unplugging of USB device during running data transfer. Possible consequences: Loss of data and damage of media possible. Remedy: Do not unplug USB device during active data transfer. (e.g. Export). Use removable media only as transfer media. Removable media are not allowed for archiving purpose. | USB device during active data transfer | 12.04% | 213 | 73 | CAUTION Cause: Risk of unplugging of USB device during running data transfer. Possible consequences: Loss of data and damage of media possible. Remedy: Do not unplug USB device during active data transfer. (e.g. Export). Use removable media only as transfer media. Removable media are not allowed for archiving purpose. | TRUE
CAUTION Cause: Software Installation will disturb system functions Possible consequences: Interruption of patient treatment and loss of patient data Remedy: Do not start installation during patient treatment | and loss of patient data | 11.79% | 248 | 220 | CAUTION Cause: Software Installation will disturb system functions Possible consequences: Interruption of patient treatment and loss of patient data Remedy: Do not start installation during patient treatment | TRUE
Which parts must be removed prior to scrapping and how these parts are disposed in an environmental friendly manner are documented in the Disposal Instructions. | in an environmental | 12.35% | 51 | 264 | Which parts must be removed prior to scrapping and how these parts are disposed in an environmental friendly manner are documented in the Disposal Instructions. | FALSE
Warning : Cause: Use of non-approved or manipulated software or hardware components Possible consequences: Risk of malfunctions posing a hazard to patients and equipment Remedy: Only use software or hardware components approved by Siemens. Perform repairs only with our expressed written approval. | ed software or hardware components | 11.33% | 329 | 1062 | Warning : Cause: Use of non-approved or manipulated software or hardware components Possible consequences: Risk of malfunctions posing a hazard to patients and equipment Remedy: Only use software or hardware components approved by Siemens. Perform repairs only with our expressed written approval. | TRUE
CAUTION Cause: Brake failure Possible consequences: Risk of crushing Remedy: Before starting the treatment, carry out the daily inspection of function and security. | efore starting the | 11.38% | 217 | 1144 | CAUTION Cause: Brake failure Possible consequences: Risk of crushing Remedy: Before starting the treatment, carry out the daily inspection of function and security. | TRUE
CAUTION Cause: Installation of the Update package failed Possible consequences: Failed Software installation will result in an undefined system state Remedy: Do not use the system and call the UPTIME Service Center | the UPTIME Service Center | 11.87% | 248 | 218 | CAUTION Cause: Installation of the Update package failed Possible consequences: Failed Software installation will result in an undefined system state Remedy: Do not use the system and call the UPTIME Service Center | TRUE
CAUTION Cause: Hot surface Possible Consequences: Risk of Burnings Remedy: While replacing trolley control panel (TUI), please note, surface on backside of TUI maybe hot. Handle with care! | trolley control panel | 11.92% | 63 | 258 | CAUTION Cause: Hot surface Possible Consequences: Risk of Burnings Remedy: While replacing trolley control panel (TUI), please note, surface on backside of TUI maybe hot. Handle with care! | FALSE
A label shall be attached near the system batteries instructing the technician to consult the service manual before exchanging the batteries. | ing the batteries | 11.56% | 305 | 201 | A label shall be attached near the system batteries instructing the technician to consult the service manual before exchanging the batteries. | FALSE
The system user does not have access to the system level. This includes the access to the underlaying operating system and the belonging functions like file menu, system properties or start menu. | e access to the system | 11.62% | 251 | 1058 | The system user does not have access to the system level. This includes the access to the underlaying operating system and the belonging functions like file menu, system properties or start menu. | FALSE
Cause: Weight of the device Possible consequences: Risk of crushing Remedy: The weight of the device can make fast braking more difficult. Pay attention to your speed as well as floor covering, slope and uneven areas. | to your speed as well as | 11.42% | 93 | 768 | Cause: Weight of the device Possible consequences: Risk of crushing Remedy: The weight of the device can make fast braking more difficult. Pay attention to your speed as well as floor covering, slope and uneven areas. | TRUE
In Service Documentation CAUTION Cause: Radiation released from a footswitch in another room Possible Consequences: Non-clinical exposure to ionizing radiation Remedy: Before performing any service activities with the wireless footswitch, check that the footswitch belongs to the system and check that the indicating label is in place and properly visible. | Non-clinical exposure to ionizing radiation | 11.81% | 303 | 353 | In Service Documentation CAUTION Cause: Radiation released from a footswitch in another room Possible Consequences: Non-clinical exposure to ionizing radiation Remedy: Before performing any service activities with the wireless footswitch, check that the footswitch belongs to the system and check that the indicating label is in place and properly visible. | FALSE
Monoblock shall keep compliance with Pressure vessels requriement from IEC60601-1 cl.9.7. | IEC60601-1 | 12.09% | 271 | 956 | Monoblock shall keep compliance with Pressure vessels requriement from IEC60601-1 cl.9.7. | FALSE
Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a heartbeat and link loss detection. | able data transfer | 12.58% | 219 | 859 | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a heartbeat and link loss detection. | FALSE
Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a heartbeat and link loss detection. | able data transfer | 12.58% | 219 | 863 | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a heartbeat and link loss detection. | FALSE
see AT-S3-28551 - Missing cross-reference and AT-S3-28552 - Missing cross-reference | reference | 12.05% | 310 | 634 | see AT-S3-28551 - Missing cross-reference and AT-S3-28552 - Missing cross-reference | FALSE
see AT-S3-28551 - Missing cross-reference and AT-S3-28552 - Missing cross-reference | reference | 12.05% | 310 | 631 | see AT-S3-28551 - Missing cross-reference and AT-S3-28552 - Missing cross-reference | FALSE
Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a heartbeat and link loss detection. | able data transfer | 12.58% | 219 | 571 | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a heartbeat and link loss detection. | FALSE
After each calibration the length of the lines in the image is updated. If no calibration has been performed, the length of the lines is displayed in pixels. | no calibration has | 12.74% | 193 | 676 | After each calibration the length of the lines in the image is updated. If no calibration has been performed, the length of the lines is displayed in pixels. | FALSE
The protection module hardware shall execute a self-test of its safety relevant hardware at startup. When the test fails, all motorized movements shall remain inhibited. | motorized movements | 12.21% | 74 | 808 | The protection module hardware shall execute a self-test of its safety relevant hardware at startup. When the test fails, all motorized movements shall remain inhibited. | FALSE
CAUTION Cause: Falling below the safety distance Possible consequences: Risk of crushing Remedy: During daily inspection of speaker: check to ensure speaker works well. | the safety distance | 11.7% | 326 | 797 | CAUTION Cause: Falling below the safety distance Possible consequences: Risk of crushing Remedy: During daily inspection of speaker: check to ensure speaker works well. | TRUE
In addition to SW a separate hardware signal shall be checked to be set 'inactive' during start-up of the x-ray generator. | In addition to | 12.1% | 194 | 632 | In addition to SW a separate hardware signal shall be checked to be set 'inactive' during start-up of the x-ray generator. | FALSE
CAUTION Cause: Braking power not sufficient. Possible consequences: Risk of crushing! Remedy: When parking the C-arm system on an inclined surface, lock the brakes on the C-arm system and turn the wheels perpendicular to the slope. | Braking power not sufficient. | 12.39% | 328 | 762 | CAUTION Cause: Braking power not sufficient. Possible consequences: Risk of crushing! Remedy: When parking the C-arm system on an inclined surface, lock the brakes on the C-arm system and turn the wheels perpendicular to the slope. | TRUE
The orbital movement has to be designed in a way that a patient can be salvaged. Note: This is needed for rescue reasons (e.g. in case of patient's unconsciousness). | orbital movement | 10.65% | 81 | 750 | The orbital movement has to be designed in a way that a patient can be salvaged. Note: This is needed for rescue reasons (e.g. in case of patient's unconsciousness). | FALSE
CAUTION Cause: Change of weight balance Possible consequences: Risk of crushing Remedy: Adding weight, detector- or monoblock-sided, means loss of weight balance and could lead to an unintentional movement of the C-arm. | al movement of the C-arm | 10.81% | 80 | 784 | CAUTION Cause: Change of weight balance Possible consequences: Risk of crushing Remedy: Adding weight, detector- or monoblock-sided, means loss of weight balance and could lead to an unintentional movement of the C-arm. | TRUE
WARNING Cause: Fluid or blood seeps into the system, use of inappropriate cleaning agents. Possible consequences: Risk of damage to the system, risk of infection or electrical hazard. Remedy: Use sterile or sterilizable covers. Do not let cleaning liquids seep into the openings of the system (e.g. air openings, gaps between covers). Observe the cleaning and disinfection instructions. Only use the recommended substances for cleaning and disinfection. | the cleaning and disinfection instructions. | 9.65% | 254 | 948 | WARNING Cause: Fluid or blood seeps into the system, use of inappropriate cleaning agents. Possible consequences: Risk of damage to the system, risk of infection or electrical hazard. Remedy: Use sterile or sterilizable covers. Do not let cleaning liquids seep into the openings of the system (e.g. air openings, gaps between covers). Observe the cleaning and disinfection instructions. Only use the recommended substances for cleaning and disinfection. | TRUE
In order to avoid hazards from fire or smoke by overheating or flammable materials the system has to be developed in accordance with the IEC 60601-1 standard. For IT subsystems it is sufficient to fullfill the IEC 60950-1. Note: Subsystem AppHost only suitable for system with it. | in accordance with the IEC | 9.61% | 234 | 1218 | In order to avoid hazards from fire or smoke by overheating or flammable materials the system has to be developed in accordance with the IEC 60601-1 standard. For IT subsystems it is sufficient to fullfill the IEC 60950-1. Note: Subsystem AppHost only suitable for system with it. | FALSE
For transfer of xray images via NXS some additional information (like patient quadrupel) shall be incorporated to identify clearly which xray image belong to which patient and the properties of the xray image. | dditional information | 10.53% | 139 | 862 | For transfer of xray images via NXS some additional information (like patient quadrupel) shall be incorporated to identify clearly which xray image belong to which patient and the properties of the xray image. | FALSE
Caution in the user manual: Cause: Risk that network connection may not be permanently available. Possible consequences: Unavailability of reference images during an intervention.. Remedy: The user should retrieve necessary images from a network archive before beginning an intervention. | network connection may not be | 10.62% | 326 | 50 | Caution in the user manual: Cause: Risk that network connection may not be permanently available. Possible consequences: Unavailability of reference images during an intervention.. Remedy: The user should retrieve necessary images from a network archive before beginning an intervention. | FALSE
CAUTION Cause: Transport damages Possible consequences: Risk of injury Remedy: Carry out professional repairs only, otherwise consequential damage to the product cannot be ruled out! | cannot be ruled out | 10.81% | 45 | 693 | CAUTION Cause: Transport damages Possible consequences: Risk of injury Remedy: Carry out professional repairs only, otherwise consequential damage to the product cannot be ruled out! | TRUE
On monitor the actual storage capacity is shown and updated when an image is stored. The image system displays the storage space still available as relative part of the assignable capacity. | storage capacity | 9.47% | 276 | 653 | On monitor the actual storage capacity is shown and updated when an image is stored. The image system displays the storage space still available as relative part of the assignable capacity. | FALSE
The overtravel (stopping distance) may be at most 10 mm. All motor-driven motions of the system are only possible by continuous and active actuation of movement switches. | of the system a | 9.41% | 269 | 791 | The overtravel (stopping distance) may be at most 10 mm. All motor-driven motions of the system are only possible by continuous and active actuation of movement switches. | FALSE
Warning Cause: Breakage or wear of the spindle nut Possible consequences: Risk of crushing Remedy: If the mechanical safeguard has had effect, the service provider must be called to replace the mechanical safeguard before the ME device can be used once more. | f the mechanical safeguard | 10.34% | 44 | 773 | Warning Cause: Breakage or wear of the spindle nut Possible consequences: Risk of crushing Remedy: If the mechanical safeguard has had effect, the service provider must be called to replace the mechanical safeguard before the ME device can be used once more. | TRUE
Usage of laser diodes of laser class 1 (according to IEC 60825-1:2014). Note: With usage of laser class 1 patient/user cannot be harmed while looking into laser light. | according to IEC 60 | 11.11% | 275 | 849 | Usage of laser diodes of laser class 1 (according to IEC 60825-1:2014). Note: With usage of laser class 1 patient/user cannot be harmed while looking into laser light. | FALSE
CAUTION Cause: Stay in emdangered area Possible consequences: Risk of unwanted radiation Remedy: Make sure that during radiation you stand in a proper location where unwanted radiation is as little as possible. | of unwanted radiation | 10.23% | 36 | 638 | CAUTION Cause: Stay in emdangered area Possible consequences: Risk of unwanted radiation Remedy: Make sure that during radiation you stand in a proper location where unwanted radiation is as little as possible. | TRUE
CAUTION Cause: Geometrical calibration misaligned Possible consequences: Risk of misjudgment Remedy: Run regular calibrations. | calibration m | 10.85% | 192 | 744 | CAUTION Cause: Geometrical calibration misaligned Possible consequences: Risk of misjudgment Remedy: Run regular calibrations. | TRUE
Supervision of power: The system shall provide a protection against excessive current and voltage. Alternatively, a test shall show that a short circuit does not result in a hazardous situation. | a hazardous situation | 11.11% | 4 | 702 | Supervision of power: The system shall provide a protection against excessive current and voltage. Alternatively, a test shall show that a short circuit does not result in a hazardous situation. | FALSE
CAUTION Cause: Image orientation mixed up Possible Consequences: Wrong interventional decision Remedy: It is recommended to use lead letters for markering the patient position. Hint: This can be included in user manual as attention note (use remedy only). | rong interventional decision | 10.69% | 156 | 681 | CAUTION Cause: Image orientation mixed up Possible Consequences: Wrong interventional decision Remedy: It is recommended to use lead letters for markering the patient position. Hint: This can be included in user manual as attention note (use remedy only). | TRUE
For transfer of xray images via NXS the data integrity of these images shall be considered by using a check sum algorithm and a ID/counter of transfered images. | be considered by | 11.25% | 9 | 860 | For transfer of xray images via NXS the data integrity of these images shall be considered by using a check sum algorithm and a ID/counter of transfered images. | FALSE
The acoustic energy / audible nosie has to be smaller than /equal to 80 dB(A). (referring to IEC60601-1, chapter 9.6.2.1) | has to be s | 9.68% | 239 | 1203 | The acoustic energy / audible nosie has to be smaller than /equal to 80 dB(A). (referring to IEC60601-1, chapter 9.6.2.1) | FALSE
CAUTION Cause: Monitor arm not locked during transportation Possible consequences: Risk of crushing Remedy: Place monitor arm in storage position before transportation. | Risk of crushing | 9.3% | 105 | 1131 | CAUTION Cause: Monitor arm not locked during transportation Possible consequences: Risk of crushing Remedy: Place monitor arm in storage position before transportation. | FALSE
Parts of the system that are suspended or supported shall be developed using the tensile safety factors listed in IEC60601-1, Table 21. These parts shall fulfill the requirements of section 9.8 of the standard. | the requirements of | 10% | 21 | 824 | Parts of the system that are suspended or supported shall be developed using the tensile safety factors listed in IEC60601-1, Table 21. These parts shall fulfill the requirements of section 9.8 of the standard. | FALSE
CAUTION Cause: Detector cover broken Possible consequences: Toxic contamination Remedy: If the detector cover is broken do not use the system at all. Please call Siemens service. | the detector cover | 10.27% | 308 | 1125 | CAUTION Cause: Detector cover broken Possible consequences: Toxic contamination Remedy: If the detector cover is broken do not use the system at all. Please call Siemens service. | FALSE
CAUTION Cause: Wrong calibration value is entered or length or angle information are not interpreted correctly Possible Consequences: Insufficient or incorrect basis for diagnostic or therapeutic decisions Remedy: Note that the measuring results of the functions Distance and Angle depend on various factors, e. g. the position of the calibration object. Note that the acquired images represent two-dimensional projections of real objects. Verify the results before diagnostic or therapeutic decisions are taken. | ote that the acquired images represent two-dimensional | 10.49% | 190 | 77 | CAUTION Cause: Wrong calibration value is entered or length or angle information are not interpreted correctly Possible Consequences: Insufficient or incorrect basis for diagnostic or therapeutic decisions Remedy: Note that the measuring results of the functions Distance and Angle depend on various factors, e. g. the position of the calibration object. Note that the acquired images represent two-dimensional projections of real objects. Verify the results before diagnostic or therapeutic decisions are taken. | TRUE
The radiation indicator circuit shall be monitored (e.g.current in indicator circuit when radiation is emitted). When an error is detected, the system shall show an appropriate error message (no blocking of radiation). | he radiation indicator | 10.45% | 129 | 1072 | The radiation indicator circuit shall be monitored (e.g.current in indicator circuit when radiation is emitted). When an error is detected, the system shall show an appropriate error message (no blocking of radiation). | FALSE
see Risk Analysis - Identified Risk Control Measures - AT-S3-27596 - User Manual - System calibration Note: VA21 is only suitable for FD version. | suitable for | 9.46% | 310 | 429 | see Risk Analysis - Identified Risk Control Measures - AT-S3-27596 - User Manual - System calibration Note: VA21 is only suitable for FD version. | FALSE
The imaging system shall store acquired 2D images immediately to the hard disk, before any user manipulation is done. The operator shall always be able to restore the original image visualization. | The imaging system s | 10.2% | 90 | 874 | The imaging system shall store acquired 2D images immediately to the hard disk, before any user manipulation is done. The operator shall always be able to restore the original image visualization. | FALSE
CAUTION Cause: Movements during patient positioning Possible Consequences: Squeezing of patient Remedy: Highlight in user manual the squeezing zones (incl. illustration). Hint: This can be included in user manual as attention note (use remedy only). | during patient positioning | 10.55% | 330 | 129 | CAUTION Cause: Movements during patient positioning Possible Consequences: Squeezing of patient Remedy: Highlight in user manual the squeezing zones (incl. illustration). Hint: This can be included in user manual as attention note (use remedy only). | TRUE
If squeeze position is exceeded during downward movement: lifting column movement will be halt; if continue to press down button, downward movement is allowed with warning sound. | lifting column move | 10.56% | 77 | 911 | If squeeze position is exceeded during downward movement: lifting column movement will be halt; if continue to press down button, downward movement is allowed with warning sound. | FALSE
CAUTION Cause: Non-harmonized configuration of HIS/RIS Possible consequences: Impossibility of Communication Remedy: Do not use configuration 'patient ID only' within HIS/RIS as Imaging System identifies a patient by quadruple (patient name, ID, sex, date-of-birth). | (patient name, ID, sex, date | 10.7% | 116 | 869 | CAUTION Cause: Non-harmonized configuration of HIS/RIS Possible consequences: Impossibility of Communication Remedy: Do not use configuration 'patient ID only' within HIS/RIS as Imaging System identifies a patient by quadruple (patient name, ID, sex, date-of-birth). | TRUE
Operator Manual CAUTION Cause: Beam of the laser light localizer strikes patient's or operator's eye. Possible consequences: Eye injury Remedy: Do not look into the beam! Take care that neither you nor any other person look directly into the light beam. | the laser light localizer | 10.51% | 103 | 496 | Operator Manual CAUTION Cause: Beam of the laser light localizer strikes patient's or operator's eye. Possible consequences: Eye injury Remedy: Do not look into the beam! Take care that neither you nor any other person look directly into the light beam. | FALSE
CAUTION Cause: Changing the grid. Possible consequences: Crushing between the housings of the detector case. Remedy: Pay attention to danger zones / danger points of crushing between the housing of the detector case. | n the housing of the | 9.5% | 253 | 509 | CAUTION Cause: Changing the grid. Possible consequences: Crushing between the housings of the detector case. Remedy: Pay attention to danger zones / danger points of crushing between the housing of the detector case. | TRUE
Labeling according to XPQR 4.8/01: 5521906 AND 715 (Labeling included, W-Lan Symbol and Symbol __TBI__please read user manual__TBI__) | according to | 10% | 320 | 1018 | Labeling according to XPQR 4.8/01: 5521906 AND 715 (Labeling included, W-Lan Symbol and Symbol __TBI__please read user manual__TBI__) | FALSE
CAUTION Cause: Rotation of c-arm Possible consequences: Risk of crushing Remedy: Pay special attention to all crushing hazards between the system's moving parts and their guide openings. | their guide openings. | 10.99% | 33 | 930 | CAUTION Cause: Rotation of c-arm Possible consequences: Risk of crushing Remedy: Pay special attention to all crushing hazards between the system's moving parts and their guide openings. | TRUE
CAUTION Cause: Improper handle of monitor arm Possible consequences: Risk of crushing Remedy: Pay special attention to all crushing hazards between the system's moving parts and their guide openings. | their guide openings. | 10.29% | 33 | 928 | CAUTION Cause: Improper handle of monitor arm Possible consequences: Risk of crushing Remedy: Pay special attention to all crushing hazards between the system's moving parts and their guide openings. | FALSE
The patient name and an additional configurable information entity (either study ID or request ID or accession number or date of birth) of the selected study shall be displayed in the patient information area on monitor A. | of the selected study s | 10.81% | 167 | 551 | The patient name and an additional configurable information entity (either study ID or request ID or accession number or date of birth) of the selected study shall be displayed in the patient information area on monitor A. | FALSE
CAUTION Cause: Inappropriate accessory. Possible consequences: The use of ACCESSORY equipment not complying with the equivalent safety requirements of this equipment may lead to a reduced level of safety of the resulting system Remedy: Use only original Siemens accessories or accessories approved by Siemens. | accessories approved by Siemens | 10.19% | 46 | 896 | CAUTION Cause: Inappropriate accessory. Possible consequences: The use of ACCESSORY equipment not complying with the equivalent safety requirements of this equipment may lead to a reduced level of safety of the resulting system Remedy: Use only original Siemens accessories or accessories approved by Siemens. | FALSE
CAUTION Cause: Rotation of c-arm Possible consequences: Risk of crushing Remedy: Pay special attention to all crushing hazards between the system's moving parts and their guide openings. | their guide openings. | 11.11% | 33 | 755 | CAUTION Cause: Rotation of c-arm Possible consequences: Risk of crushing Remedy: Pay special attention to all crushing hazards between the system's moving parts and their guide openings. | TRUE
Warning: Cause: The positions marked in illustrations to indicate danger zones of the system. Possible consequences: Danger of injury to patients or operating personnel by crushing or colliding! Remedy: Pay attention to the marked danger areas. | Pay attention to the | 8.91% | 99 | 917 | Warning: Cause: The positions marked in illustrations to indicate danger zones of the system. Possible consequences: Danger of injury to patients or operating personnel by crushing or colliding! Remedy: Pay attention to the marked danger areas. | FALSE
When the image is processed for transfer, a consistency-check is done between the study UID in the job list and the study UID in the image header. For data exchange between imaging system and RIS the DICOM3 standard is used. | imaging system and | 8.93% | 28 | 549 | When the image is processed for transfer, a consistency-check is done between the study UID in the job list and the study UID in the image header. For data exchange between imaging system and RIS the DICOM3 standard is used. | FALSE
CAUTION Cause: Wrong or extreme image processing parameters Possible consequences: Risk of limited image quality in resulting X-ray images. Incorrect interventional decision/artificial artifacts. Remedy: Apply processing parameters carefully, for example edge enhancement, windowing and always check image quality visually before performing an interventional decision. Charm: SP_00167271: correction according precondition | or extreme image processing parameters | 9.13% | 328 | 57 | CAUTION Cause: Wrong or extreme image processing parameters Possible consequences: Risk of limited image quality in resulting X-ray images. Incorrect interventional decision/artificial artifacts. Remedy: Apply processing parameters carefully, for example edge enhancement, windowing and always check image quality visually before performing an interventional decision. Charm: SP_00167271: correction according precondition | FALSE
CAUTION Cause: Incorrect collimator setting Possible Consequences: Dose not as desired Remedy: Collimator leaves have to be adjusted in a way, that at least one edge of the collimator leaves is visible in the fluoro image. This has to be checked on a daily basis. Hint: This can be included in user manual as attention note (use remedy only). | leaves is visible in the fluoro | 8.88% | 131 | 500 | CAUTION Cause: Incorrect collimator setting Possible Consequences: Dose not as desired Remedy: Collimator leaves have to be adjusted in a way, that at least one edge of the collimator leaves is visible in the fluoro image. This has to be checked on a daily basis. Hint: This can be included in user manual as attention note (use remedy only). | TRUE
If no more storage capacity is available on imaging system (i.e. the max. amount of images due to license storage options is reached during a running acquisition), acquisition is not stopped. The user is informed that no longer storage of images is possible. Still acquisition without storage is possible. | acquisition is not stopped | 8.85% | 329 | 877 | If no more storage capacity is available on imaging system (i.e. the max. amount of images due to license storage options is reached during a running acquisition), acquisition is not stopped. The user is informed that no longer storage of images is possible. Still acquisition without storage is possible. | FALSE
Service Documentation CAUTION Cause: ESD Possible Consequences: Damage of electronic components leading to deteriorations in system functionality. Remedy: Please follow the ESD guidlines during all work steps. | Please follow the | 8.92% | 298 | 254 | Service Documentation CAUTION Cause: ESD Possible Consequences: Damage of electronic components leading to deteriorations in system functionality. Remedy: Please follow the ESD guidlines during all work steps. | FALSE
If the movement enable signal is still active after the emergency stop is deactivated again, the system shall block movements, display an error message, until a normal deactivation of the enable signal and re-activation took place. | is deactivated again | 9.01% | 234 | 801 | If the movement enable signal is still active after the emergency stop is deactivated again, the system shall block movements, display an error message, until a normal deactivation of the enable signal and re-activation took place. | FALSE
CAUTION Cause: Incorrect calibration parameters. Reduced image quality Possible consequences: Risk of unnecessary X-ray exposure due to wrong calibration parameters. Remedy: Re-calibrate the system every 24 months to maintain the image quality. | Incorrect calibration | 9.31% | 191 | 522 | CAUTION Cause: Incorrect calibration parameters. Reduced image quality Possible consequences: Risk of unnecessary X-ray exposure due to wrong calibration parameters. Remedy: Re-calibrate the system every 24 months to maintain the image quality. | FALSE
To reduce probability of unintentional loss of data, the deletion of images/series is performed by a security dialog. If unarchived images/series are selected for deletion, the user has to confirm this in addition. | ed images/series are | 9.81% | 181 | 436 | To reduce probability of unintentional loss of data, the deletion of images/series is performed by a security dialog. If unarchived images/series are selected for deletion, the user has to confirm this in addition. | FALSE
CAUTION Cause: Wrong calibration value is entered or length or angle information are not interpreted correctly Possible Consequences: Insufficient or incorrect basis for diagnostic or therapeutic decisions Remedy: Note that the measuring results of the functions Distance and Angle depend on various factors, e. g. the position of the calibration object. Note that the acquired images represent two-dimensional projections of real objects. Verify the results before diagnostic or therapeutic decisions are taken. | ote that the acquired images represent two-dimensional | 10.49% | 190 | 675 | CAUTION Cause: Wrong calibration value is entered or length or angle information are not interpreted correctly Possible Consequences: Insufficient or incorrect basis for diagnostic or therapeutic decisions Remedy: Note that the measuring results of the functions Distance and Angle depend on various factors, e. g. the position of the calibration object. Note that the acquired images represent two-dimensional projections of real objects. Verify the results before diagnostic or therapeutic decisions are taken. | TRUE
CAUTION Cause: Electromagnetic interference of the system. Possible consequences: Degradation of the image quality or accuracy of the radiation parameters! Remedy: - Remove the source of interference (for instance, a mobile phone) from the area. - Shut down the system and turn it on again. - Pay attention to error messages during restart. | Pay attention to error messages | 9.17% | 99 | 1206 | CAUTION Cause: Electromagnetic interference of the system. Possible consequences: Degradation of the image quality or accuracy of the radiation parameters! Remedy: - Remove the source of interference (for instance, a mobile phone) from the area. - Shut down the system and turn it on again. - Pay attention to error messages during restart. | FALSE
Functionality that would cause dialogs or other UI elements to cover the live image or that would cause performance bottlenecks for live image processing and display are disabled during radiation. | e image processing | 9.69% | 328 | 1008 | Functionality that would cause dialogs or other UI elements to cover the live image or that would cause performance bottlenecks for live image processing and display are disabled during radiation. | FALSE
CAUTION Cause: Fire inside or in the vicinity of the system. Possible consequences: Risk of injury of patient and personnel and damage to the property. Risk of gas poisoning caused by burning plastic. Remedy: Switch off the system in the event of fire. Make sure that you and the patient know where the escape routes are. Make sure that you know where the fire extinguishers are located and familiarize yourself with the use of them. | e fire extinguishers are located and | 8.39% | 31 | 977 | CAUTION Cause: Fire inside or in the vicinity of the system. Possible consequences: Risk of injury of patient and personnel and damage to the property. Risk of gas poisoning caused by burning plastic. Remedy: Switch off the system in the event of fire. Make sure that you and the patient know where the escape routes are. Make sure that you know where the fire extinguishers are located and familiarize yourself with the use of them. | TRUE
When CD/DVD write is started the selected images can not be deleted before the write preparing is finished and the background writing has started. When writing is completed, the image is flagged to show the successful transfer to CD/DVD. | the selected images | 8.79% | 206 | 72 | When CD/DVD write is started the selected images can not be deleted before the write preparing is finished and the background writing has started. When writing is completed, the image is flagged to show the successful transfer to CD/DVD. | FALSE
An image display delay time of max. 200 ms is acceptable to ensure sufficient reaction time for the operator during intervention procedures for all imaging modes except Single Image. For Single Image, an image display delay time up to 4s is acceptable. | during intervention | 8.3% | 183 | 643 | An image display delay time of max. 200 ms is acceptable to ensure sufficient reaction time for the operator during intervention procedures for all imaging modes except Single Image. For Single Image, an image display delay time up to 4s is acceptable. | FALSE
CAUTION Cause: Wrong interpretation of results Possible Consequences: Wrong interventional decision and unnecessary stress on tissue and/or organs by shock-wave application Remedy: While using lithotripter and Cios C-arm system in combined mode, please do not perform image manipulations during therapy. | rong interventional decision | 9.09% | 156 | 668 | CAUTION Cause: Wrong interpretation of results Possible Consequences: Wrong interventional decision and unnecessary stress on tissue and/or organs by shock-wave application Remedy: While using lithotripter and Cios C-arm system in combined mode, please do not perform image manipulations during therapy. | FALSE
CAUTION Cause: LED indicator fails Possible consequences: No treatment possible Remedy: Before starting the treatment, carry out the daily inspection of function of the LED indicator. Please call service technician if a LED indicator has failed. | of the LED indicator | 8.4% | 26 | 123 | CAUTION Cause: LED indicator fails Possible consequences: No treatment possible Remedy: Before starting the treatment, carry out the daily inspection of function of the LED indicator. Please call service technician if a LED indicator has failed. | TRUE
Cleaning of the monitor is covered in cleaning guide. Reference to Caution to follow the cleaning instruction , Same as AT-S3-28683 - Missing cross-reference. The monitor must withstand the procedure. Regarding DIN 6868-157: 5 c) Before starting these tests, the viewing surface of the IMAGE DISPLAY DEVICE is to be cleaned in accordance with operating instructions, if needed. | follow the cleaning instruction | 8.29% | 253 | 598 | Cleaning of the monitor is covered in cleaning guide. Reference to Caution to follow the cleaning instruction , Same as AT-S3-28683 - Missing cross-reference. The monitor must withstand the procedure. Regarding DIN 6868-157: 5 c) Before starting these tests, the viewing surface of the IMAGE DISPLAY DEVICE is to be cleaned in accordance with operating instructions, if needed. | FALSE
CAUTION Cause: Charging capacity worn out. Possible consequences: Risk of delayed or no diagnosis. Remedy: Please regard however that the life of an Uninterruptible Power Supply is limited. | Uninterruptible | 8.85% | 56 | 1047 | CAUTION Cause: Charging capacity worn out. Possible consequences: Risk of delayed or no diagnosis. Remedy: Please regard however that the life of an Uninterruptible Power Supply is limited. | TRUE
CAUTION Cause: Illegal application of the system beyond the intended use Possible consequences: Danger for patient or user Remedy: The system is only allowed to be used within the defined intended use of clinical applications | of the system beyond | 9.57% | 325 | 1012 | CAUTION Cause: Illegal application of the system beyond the intended use Possible consequences: Danger for patient or user Remedy: The system is only allowed to be used within the defined intended use of clinical applications | TRUE
For radiography, the storage rate shall be not be changeable. Every frame shall be stored. Replay of scenes shall apply the same behaviour as live (fps). The storage rate in fluoroscopy may be reduced by the customer. | in fluoroscopy m | 7.8% | 264 | 650 | For radiography, the storage rate shall be not be changeable. Every frame shall be stored. Replay of scenes shall apply the same behaviour as live (fps). The storage rate in fluoroscopy may be reduced by the customer. | FALSE
For transfer of frames via NGS the data integrity of these frames shall be considered by using a check sum algorithm (e.g CRC32). In case of an incorrect checksum is detected an audible or visual indication shall be done. | be considered by | 8.14% | 9 | 568 | For transfer of frames via NGS the data integrity of these frames shall be considered by using a check sum algorithm (e.g CRC32). In case of an incorrect checksum is detected an audible or visual indication shall be done. | FALSE
For static images, the grey scale value of a pixel may vary by +/- 2% of the original grey scale value. For live images, the image quality may be further reduced, keeping the frame rate. The LIH shall achieve the above mentioned threshold for static images again. | e image quality may be | 8.68% | 158 | 861 | For static images, the grey scale value of a pixel may vary by +/- 2% of the original grey scale value. For live images, the image quality may be further reduced, keeping the frame rate. The LIH shall achieve the above mentioned threshold for static images again. | FALSE
CAUTION Cause: Defect of system components. Possible consequences: Any defect of a component in the system may lead to a reduced level of safety of the resulting system Remedy: Terminate the operation of the system. Notify the service provider. | operation of the system | 9.6% | 19 | 894 | CAUTION Cause: Defect of system components. Possible consequences: Any defect of a component in the system may lead to a reduced level of safety of the resulting system Remedy: Terminate the operation of the system. Notify the service provider. | TRUE
CATION Cause: Using the imaging system’s hard disk as long-term archive. Possible consequences: Loss of image data Remedy: Do not use the system’s hard disk as a long-term archive! Observe local regulations for archiving X-ray images and consider usage of DICOM archive nodes. | e of DICOM archive nodes. | 8.87% | 209 | 70 | CATION Cause: Using the imaging system’s hard disk as long-term archive. Possible consequences: Loss of image data Remedy: Do not use the system’s hard disk as a long-term archive! Observe local regulations for archiving X-ray images and consider usage of DICOM archive nodes. | TRUE
If Emergency-STOP button is pressed, radiation shall be stopped, but not inhibited. When the operator activates radiation again, the system shall release radiation again Remark: The MCU monitors the status of the emergency stop and provides this information to the system (XCU). | STOP button is pressed | 7.8% | 77 | 1080 | If Emergency-STOP button is pressed, radiation shall be stopped, but not inhibited. When the operator activates radiation again, the system shall release radiation again Remark: The MCU monitors the status of the emergency stop and provides this information to the system (XCU). | FALSE
Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a update rate detection between two frames. In case of exceeding a configurable time an audible or visual indication shall be done. | able data transfer | 7.72% | 219 | 569 | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a update rate detection between two frames. In case of exceeding a configurable time an audible or visual indication shall be done. | FALSE
When reaching the safety distance (at least 25mm) while moving downwards the vertical lift is automatically stopped. Note: This measure is mitigation crushing zone located between HZE and top of chassis (next to lifting column). | the safety distance | 8.62% | 326 | 799 | When reaching the safety distance (at least 25mm) while moving downwards the vertical lift is automatically stopped. Note: This measure is mitigation crushing zone located between HZE and top of chassis (next to lifting column). | FALSE
CAUTION Cause: Braking power not sufficient. Possible consequences: Risk of crushing! Remedy: - When parking the monitor trolley on an inclined surface, face the front toward the incline and lock the brakes. - Make sure the monitor trolley is not parked on a wet floor or other places where the adhesion of the wheels is greatly reduced. | Braking power not sufficient. | 8.43% | 328 | 1244 | CAUTION Cause: Braking power not sufficient. Possible consequences: Risk of crushing! Remedy: - When parking the monitor trolley on an inclined surface, face the front toward the incline and lock the brakes. - Make sure the monitor trolley is not parked on a wet floor or other places where the adhesion of the wheels is greatly reduced. | TRUE
Storage commitment: The archive flag ('A') is set only after a storage commitment has been received. Note: This is valid in case the DICOM node is capable of sending a storage commitment. The flag is set on study and image level. | storage commitment | 8.62% | 230 | 435 | Storage commitment: The archive flag ('A') is set only after a storage commitment has been received. Note: This is valid in case the DICOM node is capable of sending a storage commitment. The flag is set on study and image level. | FALSE
For transfer of xray images via NXS some additioinal information (like patient quadrupel) shall be incorporated to identify clearly which xray image belong to which patient and the properties of the xray image. | nal information | 7.62% | 139 | 858 | For transfer of xray images via NXS some additioinal information (like patient quadrupel) shall be incorporated to identify clearly which xray image belong to which patient and the properties of the xray image. | FALSE
The trolley must be designed in a way that prevents tilting of the trolley when driving it over inclines / ramps. Remark: Compliance with DIN EN 60601-1 tilt test | ance with DIN | 8.54% | 277 | 1272 | The trolley must be designed in a way that prevents tilting of the trolley when driving it over inclines / ramps. Remark: Compliance with DIN EN 60601-1 tilt test | FALSE
CATION Cause: Risk of plugging/unplugging USB devices during acquisition Possible consequences: Behavior can make the system unstable which may affect other processes like running acquisition tasks. Loss of data, acquisition process disturbed. Remedy: Do not plug in or unplug USB devices during acquisition tasks or other critical processes. | ing USB devices during | 6.61% | 328 | 91 | CATION Cause: Risk of plugging/unplugging USB devices during acquisition Possible consequences: Behavior can make the system unstable which may affect other processes like running acquisition tasks. Loss of data, acquisition process disturbed. Remedy: Do not plug in or unplug USB devices during acquisition tasks or other critical processes. | TRUE
Service Documentation CAUTION Cause: Undetected leakage of the monoblock or cooling unit (cooling liquid, oil) Possible consequences: Functional disturbances, leaking of liquids from the system Remedy: Check the monoblock and the cooling unit for leakage of liquids during service activities and replace defective parts. | Functional disturbances | 7.06% | 10 | 104 | Service Documentation CAUTION Cause: Undetected leakage of the monoblock or cooling unit (cooling liquid, oil) Possible consequences: Functional disturbances, leaking of liquids from the system Remedy: Check the monoblock and the cooling unit for leakage of liquids during service activities and replace defective parts. | FALSE
Caution in the user manual Cause: Modifications or interference with the imaging system without authorization Possible consequences: Incorrect or no functionality of the imaging system Remedy: Do not modify or interfere with the imaging system in any way without written authorization from Siemens Healthcare (supplier). This includes replacement of hardware and installation or execution of additional software | with the imaging system i | 6.25% | 105 | 52 | Caution in the user manual Cause: Modifications or interference with the imaging system without authorization Possible consequences: Incorrect or no functionality of the imaging system Remedy: Do not modify or interfere with the imaging system in any way without written authorization from Siemens Healthcare (supplier). This includes replacement of hardware and installation or execution of additional software | FALSE
The moving positioning units must be designed so that a overload due to a collision can be handled. Remark: The component parts shall be carrying a overload in compliance with DIN EN 60601-1 (2006). Kap. 9.8, Table 21. | ance with DIN | 6.31% | 277 | 779 | The moving positioning units must be designed so that a overload due to a collision can be handled. Remark: The component parts shall be carrying a overload in compliance with DIN EN 60601-1 (2006). Kap. 9.8, Table 21. | FALSE
CAUTION Cause: Wrong focus position Possible Consequences: Wrong focus position and unnecessary stress on tissue and/or organs by shock-wave application Remedy: Before using lithotripter and Cios C-arm system in combined mode, please make sure that the X-ray central beam must run through the focus phantom in the 0° and oblique projection of the X-ray C-arm. | lease make sure that the | 6.87% | 32 | 669 | CAUTION Cause: Wrong focus position Possible Consequences: Wrong focus position and unnecessary stress on tissue and/or organs by shock-wave application Remedy: Before using lithotripter and Cios C-arm system in combined mode, please make sure that the X-ray central beam must run through the focus phantom in the 0° and oblique projection of the X-ray C-arm. | FALSE
The system is designed in a way that prevents tipping over. Remark: During a step shock test the device shall not tilt in any position and no mechanical fixings that could lead to a hazard must loosen. | The system is | 6.9% | 329 | 1162 | The system is designed in a way that prevents tipping over. Remark: During a step shock test the device shall not tilt in any position and no mechanical fixings that could lead to a hazard must loosen. | FALSE
The drive of the vertical movement must be designed in a way that prevents a autonoums movement, e.g. to provide a brake or self-locking of the drive. Remark: No autonomous movement is given, as long as a lowering within one hour does not exceed 2 mm. | ertical movement | 6.67% | 63 | 788 | The drive of the vertical movement must be designed in a way that prevents a autonoums movement, e.g. to provide a brake or self-locking of the drive. Remark: No autonomous movement is given, as long as a lowering within one hour does not exceed 2 mm. | FALSE
CAUTION Cause: Brake failure Possible Consequences: Crushing due to autonomous movements Remedy: Before starting the treatment, carry out the daily inspection of function and security. Hint: This can be included in user manual as attention note (use remedy only). | efore starting the | 7.04% | 217 | 765 | CAUTION Cause: Brake failure Possible Consequences: Crushing due to autonomous movements Remedy: Before starting the treatment, carry out the daily inspection of function and security. Hint: This can be included in user manual as attention note (use remedy only). | TRUE
Caution in the user manual: Cause: Risk of plugging/unplugging USB devices during acquisition Possible consequences: Behavior can make the system unstable which may affect other processes like running acquisition tasks. Loss of data, acquisition process disturbed. Remedy: Do not plug in or unplug USB devices during acquisition tasks or other critical processes. | ing USB devices during | 6.23% | 328 | 54 | Caution in the user manual: Cause: Risk of plugging/unplugging USB devices during acquisition Possible consequences: Behavior can make the system unstable which may affect other processes like running acquisition tasks. Loss of data, acquisition process disturbed. Remedy: Do not plug in or unplug USB devices during acquisition tasks or other critical processes. | FALSE
Caution: Cause: Loss of image data Possible Consequences: Risk of unnecessary X-ray exposure Remedy: The imaging system's hard disk is not suited as long-term archive for patient or image data. For archiving image data there are legal regulations concerning the archiving duration and availability of data, data security (data integrity) and recommendations concerning fire protection and other damages. The operator of the archive is responsible for observing these regulations. Caused by the technical progress and the legal regulations, storage and retrieval of image data cannot be realized with a single storage and media technology. Hence migration of data can become necessary in responsibility of the operator of the digital archive. Hint: This can be included in user manual as attention note (use remedy only). | archive is responsible for observing these regulations. | 6.65% | 209 | 1054 | Caution: Cause: Loss of image data Possible Consequences: Risk of unnecessary X-ray exposure Remedy: The imaging system's hard disk is not suited as long-term archive for patient or image data. For archiving image data there are legal regulations concerning the archiving duration and availability of data, data security (data integrity) and recommendations concerning fire protection and other damages. The operator of the archive is responsible for observing these regulations. Caused by the technical progress and the legal regulations, storage and retrieval of image data cannot be realized with a single storage and media technology. Hence migration of data can become necessary in responsibility of the operator of the digital archive. Hint: This can be included in user manual as attention note (use remedy only). | TRUE
CAUTION Cause: Operation of the system by non-trained users Possible consequences: Misinterpretation of image information or of indications of the system, leading to inappropriate decisions in the clinical workflow, incorrect diagnosis or treatment, delay of the clinical procedure, or exposure to radiation without clinical benefit. Remedy: The system must only be used by persons with the necessary knowledge, after an appropriate application training. | persons with the necessary | 5.91% | 9 | 1233 | CAUTION Cause: Operation of the system by non-trained users Possible consequences: Misinterpretation of image information or of indications of the system, leading to inappropriate decisions in the clinical workflow, incorrect diagnosis or treatment, delay of the clinical procedure, or exposure to radiation without clinical benefit. Remedy: The system must only be used by persons with the necessary knowledge, after an appropriate application training. | TRUE
Back up nut integrated in the ball screw nut to prevent the Spindle from collapsing if main nut would fail. A failed spindle leads normally to a not moving spindle and a test report (or third party certification) proof this case of failure. Therefore no additional indication for operator is needed. | integrated in the | 6.33% | 310 | 775 | Back up nut integrated in the ball screw nut to prevent the Spindle from collapsing if main nut would fail. A failed spindle leads normally to a not moving spindle and a test report (or third party certification) proof this case of failure. Therefore no additional indication for operator is needed. | FALSE
The relevant Service Documentation CAUTION Cause: Technician looks into the laser of the optical link (I.I side/ Tube side/Detector). Possible consequences: Risk of Eye injury. Remedy: Do not look into the laser diodes of the optical link | into the laser | 6.56% | 309 | 952 | The relevant Service Documentation CAUTION Cause: Technician looks into the laser of the optical link (I.I side/ Tube side/Detector). Possible consequences: Risk of Eye injury. Remedy: Do not look into the laser diodes of the optical link | FALSE
Accuracy of collimator leaves position shall remain unchanged during examination. It is possible to perform a search run to calibrate position of collimator leaves. In case a search run will be implemented, it has to be ensured, that no search run will be triggered and/or performed during acquisition (i.e. while releasing X-ray). | examination. It is | 5.93% | 147 | 536 | Accuracy of collimator leaves position shall remain unchanged during examination. It is possible to perform a search run to calibrate position of collimator leaves. In case a search run will be implemented, it has to be ensured, that no search run will be triggered and/or performed during acquisition (i.e. while releasing X-ray). | FALSE
The X-ray protection during manufacturing is controlled by following up 5521922 AND 507 Ionizing Radiation Protection Work Instruction. The risks in the manufacturing process are sufficiently handled by 5521906 AND 719 Process Validation and Verification Procedure. | ing is controlled | 6.39% | 67 | 151 | The X-ray protection during manufacturing is controlled by following up 5521922 AND 507 Ionizing Radiation Protection Work Instruction. The risks in the manufacturing process are sufficiently handled by 5521906 AND 719 Process Validation and Verification Procedure. | FALSE
CATION Cause: Modifications or interference with the imaging system without authorization Possible consequences: Incorrect or no functionality of the imaging system Remedy: Do not modify or interfere with the imaging system in any way without written authorization from Siemens Healthcare (supplier). This includes replacement of hardware and installation or execution of additional software | with the imaging system i | 6.57% | 105 | 89 | CATION Cause: Modifications or interference with the imaging system without authorization Possible consequences: Incorrect or no functionality of the imaging system Remedy: Do not modify or interfere with the imaging system in any way without written authorization from Siemens Healthcare (supplier). This includes replacement of hardware and installation or execution of additional software | TRUE
All buttons intended for the control of a mechanical movement of the system/parts of the system must trigger a cyclic repetition of the activation signal (only valid for the SW based signal path) to allow the system control the detection of a lost telegram and to stop the movement immediately in this case (“eigensicherer Zustand”). | ical movement of the | 6.31% | 63 | 125 | All buttons intended for the control of a mechanical movement of the system/parts of the system must trigger a cyclic repetition of the activation signal (only valid for the SW based signal path) to allow the system control the detection of a lost telegram and to stop the movement immediately in this case (“eigensicherer Zustand”). | FALSE
The system shall be developed such that the dielectric strength of electrical insulation shall be able to withstand the test voltages specified in IEC60601-1, Table 6. The test conditions are specified in section 8.8.3 of the standard. | conditions are | 6.81% | 193 | 825 | The system shall be developed such that the dielectric strength of electrical insulation shall be able to withstand the test voltages specified in IEC60601-1, Table 6. The test conditions are specified in section 8.8.3 of the standard. | FALSE
The enclosure shall be designed in a way that it can be easily cleaned and wipe disinfected. The enclosure must be compliant with IEC60601-1 chapter 11.6.6 & 7.9.2.12 regarding cleaning and disinfection requirements. Fit criterion: All different enclosure surfaces must be compliant with MFR 890/009e - Resistance of surfaces to disinfection and cleaning agents. | cleaning and disinfection | 7.32% | 254 | 160 | The enclosure shall be designed in a way that it can be easily cleaned and wipe disinfected. The enclosure must be compliant with IEC60601-1 chapter 11.6.6 & 7.9.2.12 regarding cleaning and disinfection requirements. Fit criterion: All different enclosure surfaces must be compliant with MFR 890/009e - Resistance of surfaces to disinfection and cleaning agents. | FALSE
CAUTION Cause: Insufficient charge of UPS Possible consequences: Risk of delayed or no diagnosis Remedy: The user shall check the state of charge of UPS and in case of low battery charge, the trolley is not to be unplugged from power if it is switched on. If the user wants to move the trolley, a manual shutdown has to be performed before unplugging the trolley from power. | has to be performed | 5.57% | 31 | 1049 | CAUTION Cause: Insufficient charge of UPS Possible consequences: Risk of delayed or no diagnosis Remedy: The user shall check the state of charge of UPS and in case of low battery charge, the trolley is not to be unplugged from power if it is switched on. If the user wants to move the trolley, a manual shutdown has to be performed before unplugging the trolley from power. | TRUE
Operator Manual CAUTION Cause: New viruses will not be recognized Possible consequences: System malfunction because of virus infection Remedy: Confirm virus pattern file and scan engine updates at start-up time. Only available with remote update handling. Otherwise contact Siemens Service. | se contact Siemens | 6.46% | 78 | 278 | Operator Manual CAUTION Cause: New viruses will not be recognized Possible consequences: System malfunction because of virus infection Remedy: Confirm virus pattern file and scan engine updates at start-up time. Only available with remote update handling. Otherwise contact Siemens Service. | FALSE
Motorized movement shall be possible only when two redundant signals provided by two independent paths are present (continuous activation), preventing unintended movement activation. When only one of two signals is present, all motorized movements shall be stopped and inhibited for the respective control. | preventing unintended | 7.42% | 301 | 836 | Motorized movement shall be possible only when two redundant signals provided by two independent paths are present (continuous activation), preventing unintended movement activation. When only one of two signals is present, all motorized movements shall be stopped and inhibited for the respective control. | FALSE
CAUTION Cause: Imaging system crash / freeze without recovery possibility (System shutdown has no effect to imaging system due to UPS) Possible consequences: Risk of injury for the patient Remedy: please follow the reset procedure | lease follow the | 7.23% | 298 | 1009 | CAUTION Cause: Imaging system crash / freeze without recovery possibility (System shutdown has no effect to imaging system due to UPS) Possible consequences: Risk of injury for the patient Remedy: please follow the reset procedure | TRUE
Service Documents WARNING Cause: Non-compliance to radiation safety guidelines Possible consequences: Increased probability of deterministic and stochastic radiation effects Remedy: - Follow the radiation safety guidelines and the local regulations for radiation safety - Increase the distance to the radiation source - Release radiation only when required - Use low radiation parameters - Use radiation protection, e. g. protection clothing | eterministic and stochastic | 6.19% | 9 | 271 | Service Documents WARNING Cause: Non-compliance to radiation safety guidelines Possible consequences: Increased probability of deterministic and stochastic radiation effects Remedy: - Follow the radiation safety guidelines and the local regulations for radiation safety - Increase the distance to the radiation source - Release radiation only when required - Use low radiation parameters - Use radiation protection, e. g. protection clothing | FALSE
The protection module hardware shall have a safe power supply: The over- / undervoltage protection shall be either redundantly implemented or be tested at startup. The voltage limits shall be defined according to the guaranteed working voltages of the hardware to ensure correct functionality. | ed according to the | 6.44% | 133 | 807 | The protection module hardware shall have a safe power supply: The over- / undervoltage protection shall be either redundantly implemented or be tested at startup. The voltage limits shall be defined according to the guaranteed working voltages of the hardware to ensure correct functionality. | FALSE
Caution: Cause: Front cover touches the screw stem of the transformer and the metal parts of the Main unit Possible consequences: Risk of injury of service person and damage to the property Remedy: Any parts, don’t touch the screw stem of the transformer during service | and damage to the | 6.93% | 213 | 972 | Caution: Cause: Front cover touches the screw stem of the transformer and the metal parts of the Main unit Possible consequences: Risk of injury of service person and damage to the property Remedy: Any parts, don’t touch the screw stem of the transformer during service | FALSE
Movement activation through dead man functions of the push-buttons. All buttons with dead man functionality available at the control panels must have a redundant mechanism (continuous activation) to bring buttons to their initial position after the user has stopped pressing them | the control panels | 7.14% | 149 | 810 | Movement activation through dead man functions of the push-buttons. All buttons with dead man functionality available at the control panels must have a redundant mechanism (continuous activation) to bring buttons to their initial position after the user has stopped pressing them | FALSE
Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a latency detection. This means measuring the time between sending a frame till receiving this frame and in case of exceeding a configurable limit to indicate it with an audible or visual indicator. | able data transfer | 6.07% | 219 | 570 | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a latency detection. This means measuring the time between sending a frame till receiving this frame and in case of exceeding a configurable limit to indicate it with an audible or visual indicator. | FALSE
There shall be only restricted access to internet possible for a clinical user (e.g. only some apphost related vendor URLs shall be allowed to access, but a normal internet usage shall not be possible). The service user shall be able to access internet via a special service environment | service environment | 6.97% | 248 | 373 | There shall be only restricted access to internet possible for a clinical user (e.g. only some apphost related vendor URLs shall be allowed to access, but a normal internet usage shall not be possible). The service user shall be able to access internet via a special service environment | FALSE
Design of C-arm shall be weight balanced to prevent autonomous movement with brakes open. Note: A remaining movement may be observable as weight of C-arm cable will pull C-arm in retrograde direction. Range of remaining movement depends on angular/orbital position and shall not exceed 23 degree. Nonetheless, remaining movement shall be uncritical for the user. | ement depends on | 4.67% | 193 | 786 | Design of C-arm shall be weight balanced to prevent autonomous movement with brakes open. Note: A remaining movement may be observable as weight of C-arm cable will pull C-arm in retrograde direction. Range of remaining movement depends on angular/orbital position and shall not exceed 23 degree. Nonetheless, remaining movement shall be uncritical for the user. | FALSE
Non-compliance can lead to severe bodily injury or even death, as well as property damage. After the cover panels are opened, parts under voltage are accessible. To avoid danger, disconnect the system from the power supply before opening the covers. An uninterruptible power supply(UPS) is installed in the system. Unplug the power input connector of the imaging system or switch off the UPS prior to beginning service work. If steps have to be performed using electrical power, the general safety information in 'Medical Product; Safety Inforamtion; General Safety Notes' must be observed. | n uninterruptible power supply | 5.06% | 98 | 250 | Non-compliance can lead to severe bodily injury or even death, as well as property damage. After the cover panels are opened, parts under voltage are accessible. To avoid danger, disconnect the system from the power supply before opening the covers. An uninterruptible power supply(UPS) is installed in the system. Unplug the power input connector of the imaging system or switch off the UPS prior to beginning service work. If steps have to be performed using electrical power, the general safety information in 'Medical Product; Safety Inforamtion; General Safety Notes' must be observed. | FALSE
The trolley must be designed in a way that prevents tilting of the trolley when hitting a door frame. This is in accordance to IEC60601-1 clause 15.3.5 (Rough Handling Test). Repeat door frame test three times with speed of 0.8 m/s (tolerance: +/- 0.1 m/s). Remark: Execute door frame shock test according to standard, possibly emerged damages have to be examined for hazard (e.g. access to electrical and moveable parts has to be prevented). | in accordance to IEC | 4.7% | 266 | 1265 | The trolley must be designed in a way that prevents tilting of the trolley when hitting a door frame. This is in accordance to IEC60601-1 clause 15.3.5 (Rough Handling Test). Repeat door frame test three times with speed of 0.8 m/s (tolerance: +/- 0.1 m/s). Remark: Execute door frame shock test according to standard, possibly emerged damages have to be examined for hazard (e.g. access to electrical and moveable parts has to be prevented). | FALSE
The trolley must be designed in a way that prevents tilting of the trolley when driving it over a 40 mm step (either ascending or descending) and that prevents loosening of mechanical fixings. This is in accordance to IEC60601-1, clause 15.3.5 (Rough Handling Test). Remark: Execute step shock test according to standard, height of the step 40 mm, possibly emerged damages have to be examined for hazard (e.g. access to electrical and moveable parts has to be prevented) | in accordance to IEC | 4.43% | 266 | 1267 | The trolley must be designed in a way that prevents tilting of the trolley when driving it over a 40 mm step (either ascending or descending) and that prevents loosening of mechanical fixings. This is in accordance to IEC60601-1, clause 15.3.5 (Rough Handling Test). Remark: Execute step shock test according to standard, height of the step 40 mm, possibly emerged damages have to be examined for hazard (e.g. access to electrical and moveable parts has to be prevented) | FALSE
CAUTION Cause: Acoustic signals within system Possible consequences: Indication on possible risks, e.g. of crushing, long-time X-ray, high dose level Remedy: Please note, if system is used in same way after acoustic signal was emitted, described risks may occur. Use system very carefully to prevent these hazards.Hint: This can be included in user manual as attention note (use remedy only). | system is used in | 4.55% | 140 | 795 | CAUTION Cause: Acoustic signals within system Possible consequences: Indication on possible risks, e.g. of crushing, long-time X-ray, high dose level Remedy: Please note, if system is used in same way after acoustic signal was emitted, described risks may occur. Use system very carefully to prevent these hazards.Hint: This can be included in user manual as attention note (use remedy only). | TRUE
The displayed live image is not covered by dialogs or other UI elements except from intentionally selected UI overlays by user interaction. Note: Intentionally selected UI overlays: Proactive display of e.g. the menu for kV-Lock at the bottom of the live image is desired by the user and therefore the user can adjust to this. | not covered by | 4.79% | 320 | 662 | The displayed live image is not covered by dialogs or other UI elements except from intentionally selected UI overlays by user interaction. Note: Intentionally selected UI overlays: Proactive display of e.g. the menu for kV-Lock at the bottom of the live image is desired by the user and therefore the user can adjust to this. | FALSE
Only selected service functionality shall be available for remote service in limited access mode. No customer functions shall be available during remote service in full access mode. Remote service shall be only possible when approved by the operator. An indication shall be provided to the user that remote service with full access has been enabled. The user shall still able to release radiation. . | to release radiation | 5.21% | 86 | 225 | Only selected service functionality shall be available for remote service in limited access mode. No customer functions shall be available during remote service in full access mode. Remote service shall be only possible when approved by the operator. An indication shall be provided to the user that remote service with full access has been enabled. The user shall still able to release radiation. . | FALSE
For all surfaces which can be touched by the patient (applied parts according to IEC 60601-1 cl. 4.6,Applied Parts and Accessible Parts), the surface temperatures under normal and single fault conditions according to IEC 60601-1, table 24 for the following contact duration shall be applied: Affected components and areas: - Cover of the detector: 10 min <= t - Cover of the monoblock, excluded monoblock backside: 1 min <= t <10 min | according to IEC 60601-1 | 5.59% | 275 | 966 | For all surfaces which can be touched by the patient (applied parts according to IEC 60601-1 cl. 4.6,Applied Parts and Accessible Parts), the surface temperatures under normal and single fault conditions according to IEC 60601-1, table 24 for the following contact duration shall be applied: Affected components and areas: - Cover of the detector: 10 min <= t - Cover of the monoblock, excluded monoblock backside: 1 min <= t <10 min | FALSE
The chassis must be designed in a way that prevents damage of components when driving over a 40 mm step (either ascending or descending). This is in accordance to IEC60601-1, clause 15.3.5 (Rough Handling Test). Remark: Execute step shock test according to standard, height of the step 40 mm, possibly emerged damages have to be examined for hazard (e.g. access to electrical and moveable parts has to be prevented). | in accordance to IEC | 5.02% | 266 | 1139 | The chassis must be designed in a way that prevents damage of components when driving over a 40 mm step (either ascending or descending). This is in accordance to IEC60601-1, clause 15.3.5 (Rough Handling Test). Remark: Execute step shock test according to standard, height of the step 40 mm, possibly emerged damages have to be examined for hazard (e.g. access to electrical and moveable parts has to be prevented). | FALSE
All power cables have to be protected against cable damage to prevent electrical shock. In addition, it has to ensured, that system's connection to power line remains active as long as possible (a connection via socket outlet and plug is not sufficient). All power cables need to be secured with cable anchorage / strain-relief. Pull forces for cable retention and cable anchorage shall be in accordance to IEC60601-1 8.11.3.5. | in accordance to IEC | 4.91% | 266 | 1117 | All power cables have to be protected against cable damage to prevent electrical shock. In addition, it has to ensured, that system's connection to power line remains active as long as possible (a connection via socket outlet and plug is not sufficient). All power cables need to be secured with cable anchorage / strain-relief. Pull forces for cable retention and cable anchorage shall be in accordance to IEC60601-1 8.11.3.5. | FALSE
Processes employed throughout the product lifetime like design, purchasing, production, modification, packing, transport, installation, service, refurbishing and disposal can be the source of failures or hazards when not carried out appropriately Currently there is no refurbishing process. | installation, | 5.14% | 303 | 20 | Processes employed throughout the product lifetime like design, purchasing, production, modification, packing, transport, installation, service, refurbishing and disposal can be the source of failures or hazards when not carried out appropriately Currently there is no refurbishing process. | FALSE
There shall be a mechanical safety mechanism for the lift column which prevents the C-arm from crashing down in case of a spindle failure. The mechanical safety mechanism shall be designed to fulfill a tensile safety factor >= 4. The overtravel (stopping distance from the time of engagement to the time of no further movement) shall be max. 5 mm. | distance from the | 5.44% | 37 | 777 | There shall be a mechanical safety mechanism for the lift column which prevents the C-arm from crashing down in case of a spindle failure. The mechanical safety mechanism shall be designed to fulfill a tensile safety factor >= 4. The overtravel (stopping distance from the time of engagement to the time of no further movement) shall be max. 5 mm. | FALSE
Emergency user login is possible without a password. Note: Due to PSS requirements, disk encryption is configurable: It is an option and must be activated by service on request. Consequences and boundary conditions are described in operator manual, admin guide and MDS2 statement. With these conditions, it is acceptable to restrict the effectivity of above mentioned measure. | s possible without | 5.01% | 35 | 581 | Emergency user login is possible without a password. Note: Due to PSS requirements, disk encryption is configurable: It is an option and must be activated by service on request. Consequences and boundary conditions are described in operator manual, admin guide and MDS2 statement. With these conditions, it is acceptable to restrict the effectivity of above mentioned measure. | FALSE
Processes employed throughout the product lifetime like design, purchasing, production, modification, packing, transport, installation, service, refurbishing and disposal can be the source of failures or hazards when not carried out appropriately Currently there is no refurbishing process. | installation, | 5.14% | 303 | 16 | Processes employed throughout the product lifetime like design, purchasing, production, modification, packing, transport, installation, service, refurbishing and disposal can be the source of failures or hazards when not carried out appropriately Currently there is no refurbishing process. | FALSE
The temperature inside the detector is checked quasi-continuously (i.e. short interval) by imaging system. According to the detector type a warning for the user is displayed, when first warning temperature (57°C) is reached. The warning message informs user on possible consequences regarding image quality. Nevertheless, system shall work properly. When the second warning temperature (58°C) is reached, an additional warning message is displayed to inform user on possible consequences regarding lifetime of detector. The detector will not be switched off at any time. | message is displayed | 3.84% | 242 | 958 | The temperature inside the detector is checked quasi-continuously (i.e. short interval) by imaging system. According to the detector type a warning for the user is displayed, when first warning temperature (57°C) is reached. The warning message informs user on possible consequences regarding image quality. Nevertheless, system shall work properly. When the second warning temperature (58°C) is reached, an additional warning message is displayed to inform user on possible consequences regarding lifetime of detector. The detector will not be switched off at any time. | FALSE
Which parts must be removed prior to scrapping and how these parts are disposed in an environmental friendly manner are documented in the Disposal Instructions (Service Documents). CAUTION Cause: ESU contains harmful and flammable liquid (Acetonitrile) and is not disposed correctly Possible Consequences: Risk of injury and environmental damage. Remedy: For proper disposal, send the ESU back to the manufacturer, SIEMENS. | in an environmental | 4.66% | 51 | 116 | Which parts must be removed prior to scrapping and how these parts are disposed in an environmental friendly manner are documented in the Disposal Instructions (Service Documents). CAUTION Cause: ESU contains harmful and flammable liquid (Acetonitrile) and is not disposed correctly Possible Consequences: Risk of injury and environmental damage. Remedy: For proper disposal, send the ESU back to the manufacturer, SIEMENS. | FALSE
The system shall check the status of both radiation release signal paths at startup of every control. When an active radiation release signal path is detected, radiation from the affected control (sharing one dead man signal) shall be blocked and an error message shall be displayed, until a deactivation and regular re-activation of both radiation release paths take place. | e radiation release s | 5.51% | 138 | 1081 | The system shall check the status of both radiation release signal paths at startup of every control. When an active radiation release signal path is detected, radiation from the affected control (sharing one dead man signal) shall be blocked and an error message shall be displayed, until a deactivation and regular re-activation of both radiation release paths take place. | FALSE
The system shall check the status of both movement release signal paths at startup. When an active movement release signal path is detected, movements from the affected control (sharing one enable hardware line) shall be blocked and an error message shall be displayed, until a deactivation and regular re-activation of both movement release paths takes place. | an error message | 4.71% | 208 | 804 | The system shall check the status of both movement release signal paths at startup. When an active movement release signal path is detected, movements from the affected control (sharing one enable hardware line) shall be blocked and an error message shall be displayed, until a deactivation and regular re-activation of both movement release paths takes place. | FALSE
Moving components such like cables, plugs, screws which are located in the proximity of isolators or wires may come into contact with voltages and increase the risk of fire or electric shocks. Mitigation is performed by fixing all moving wires, plugs connectors, etc. with means such like cable conduit, cable straps, plastic bridges or screws, glue, heat shrink tubes, etc. to avoid contacts caused by attrition. This measure is in accordance to IEC60601-1 chap. 8.10. Note: Movable components are affected by e.g. adjustment of monitor height or horizontal movement of C-arm (but not limited to). In addition, this shall address service workflows as well (e.g. pulling components from its defined position in system during maintenance or replacement). | come into contact with | 3.17% | 253 | 1189 | Moving components such like cables, plugs, screws which are located in the proximity of isolators or wires may come into contact with voltages and increase the risk of fire or electric shocks. Mitigation is performed by fixing all moving wires, plugs connectors, etc. with means such like cable conduit, cable straps, plastic bridges or screws, glue, heat shrink tubes, etc. to avoid contacts caused by attrition. This measure is in accordance to IEC60601-1 chap. 8.10. Note: Movable components are affected by e.g. adjustment of monitor height or horizontal movement of C-arm (but not limited to). In addition, this shall address service workflows as well (e.g. pulling components from its defined position in system during maintenance or replacement). | FALSE
The portable control elements except Hand-Helds with masses <= 10 kg have to be designed so that they can be dropped from a height of 5 cm without having damages that cause unacceptable risks. Hand-Helds have to be designed to withstand drop from heights at which the ME EQUIPMENT is used (as specified in the ACCOMPANYING DOCUMENTS), or from a height of 1 m, whichever is greater, onto a 50 mm +/-5 mm thick hardwood board (hardwood > 600 kg/m3) lying flat on a concrete or a similar rigid base. Test criterion: Execute drop tests according to standard IEC 60601-1 at heights of 5 cm for portable mobile control panels or heights of >= 1m for Hand-Helds. Possible emerged damages have to be examined for hazards (e.g. access to electrical and moveable parts has to be prevented). | e control elements | 2.42% | 62 | 1134 | The portable control elements except Hand-Helds with masses <= 10 kg have to be designed so that they can be dropped from a height of 5 cm without having damages that cause unacceptable risks. Hand-Helds have to be designed to withstand drop from heights at which the ME EQUIPMENT is used (as specified in the ACCOMPANYING DOCUMENTS), or from a height of 1 m, whichever is greater, onto a 50 mm +/-5 mm thick hardwood board (hardwood > 600 kg/m3) lying flat on a concrete or a similar rigid base. Test criterion: Execute drop tests according to standard IEC 60601-1 at heights of 5 cm for portable mobile control panels or heights of >= 1m for Hand-Helds. Possible emerged damages have to be examined for hazards (e.g. access to electrical and moveable parts has to be prevented). | FALSE
Moving components such like cables, plugs, screws which are located in the proximity of isolators or wires may come into contact with voltages and increase the risk of fire or electric shocks. Mitigation is performed by fixing all moving wires, plugs connectors, etc. with means such like cable conduit, cable straps, plastic bridges or screws, glue, heat shrink tubes, etc. to avoid contacts caused by attrition. This measure is in accordance to IEC60601-1 chap. 8.10. Note: Movable components are affected by e.g. adjustment of monitor height or horizontal movement of C-arm (but not limited to). In addition, this shall address service workflows as well (e.g. pulling components from its defined position in system during maintenance or replacement). | come into contact with | 3.17% | 253 | 1128 | Moving components such like cables, plugs, screws which are located in the proximity of isolators or wires may come into contact with voltages and increase the risk of fire or electric shocks. Mitigation is performed by fixing all moving wires, plugs connectors, etc. with means such like cable conduit, cable straps, plastic bridges or screws, glue, heat shrink tubes, etc. to avoid contacts caused by attrition. This measure is in accordance to IEC60601-1 chap. 8.10. Note: Movable components are affected by e.g. adjustment of monitor height or horizontal movement of C-arm (but not limited to). In addition, this shall address service workflows as well (e.g. pulling components from its defined position in system during maintenance or replacement). | FALSE
WARNING Even after switching off the power supply, electrical voltage is present in the system. Live-line working can cause electric shock or mechanical damage through short circuiting. Risk of death, serious physical injury or damage to property. De-energize the system. If checks or measurements require live-line working, cover energized parts to prevent unintentional contact. Perform live-line working with extreme caution. Use only measurement equipment and devices which are suited to the task. Read and observe the warning labels (attached next to conducting parts). | with extreme caution | 3.62% | 16 | 112 | WARNING Even after switching off the power supply, electrical voltage is present in the system. Live-line working can cause electric shock or mechanical damage through short circuiting. Risk of death, serious physical injury or damage to property. De-energize the system. If checks or measurements require live-line working, cover energized parts to prevent unintentional contact. Perform live-line working with extreme caution. Use only measurement equipment and devices which are suited to the task. Read and observe the warning labels (attached next to conducting parts). | FALSE
Service Documents WARNING Cause: Even after switching off the power supply, electrical voltage is present in the system. Live-line working can cause electric shock or mechanical damage through short circuiting. Possible Consequences: Risk of death, serious physical injury or damage to property. Remedy: De-energize the system. If checks or measurements require live-line working, cover energized parts to prevent unintentional contact. Perform live-line working with extreme caution. Use only measurement equipment and devices which are suited to the task. Read and observe the warning labels attached next to conducting parts. | with extreme caution | 3.28% | 16 | 113 | Service Documents WARNING Cause: Even after switching off the power supply, electrical voltage is present in the system. Live-line working can cause electric shock or mechanical damage through short circuiting. Possible Consequences: Risk of death, serious physical injury or damage to property. Remedy: De-energize the system. If checks or measurements require live-line working, cover energized parts to prevent unintentional contact. Perform live-line working with extreme caution. Use only measurement equipment and devices which are suited to the task. Read and observe the warning labels attached next to conducting parts. | FALSE
Supply Chain Safety Instruction WARNING Cause: If the ESU is not discharged completely, electrical voltage is present and is accessible when e. g. main unit covers are removed. Possible consequences: Risk of death or serious injury due to electric shock when electrified components are touched. Remedy: Wait until discharge finishes (10min if without ESU and 20min if equipped with ESU) for both ESU and generator capacitors before starting to work on relevant parts of the main unit. Observe the LED at the inverter box which is lit when the capacitors are charged and turned off nealy discharged. Don't disconnect ESU and generator before this time as well, otherwise the discharging circuit will not work. | death or serious injury | 3.35% | 4 | 107 | Supply Chain Safety Instruction WARNING Cause: If the ESU is not discharged completely, electrical voltage is present and is accessible when e. g. main unit covers are removed. Possible consequences: Risk of death or serious injury due to electric shock when electrified components are touched. Remedy: Wait until discharge finishes (10min if without ESU and 20min if equipped with ESU) for both ESU and generator capacitors before starting to work on relevant parts of the main unit. Observe the LED at the inverter box which is lit when the capacitors are charged and turned off nealy discharged. Don't disconnect ESU and generator before this time as well, otherwise the discharging circuit will not work. | FALSE
Make note in the relevant factory documents. How these parts are disposed in an environmentally friendly manner has to be documented in the disposal instructions. Cause: Toxic materials Possible consequences: Toxic contamination in case of detector panel is broken by mechanical shock Remedy: Handle detector with care! If the detector is broken do not use it, inspect if any particles are spilled. Leaked particles have to be collected and kept in a sealed container. Spilled particles have to be returned to Siemens. | in an environmentally | 4.22% | 51 | 1124 | Make note in the relevant factory documents. How these parts are disposed in an environmentally friendly manner has to be documented in the disposal instructions. Cause: Toxic materials Possible consequences: Toxic contamination in case of detector panel is broken by mechanical shock Remedy: Handle detector with care! If the detector is broken do not use it, inspect if any particles are spilled. Leaked particles have to be collected and kept in a sealed container. Spilled particles have to be returned to Siemens. | FALSE
Make note in the relevant service documents. How these parts are disposed in an environmentally friendly manner has to be documented in the disposal instructions. Cause: Toxic materials Possible consequences: Toxic contamination in case of detector panel is broken by mechanical shock Remedy: Handle detector with care! If the detector is broken do not use it, inspect if any particles are spilled. Leaked particles have to be collected and kept in a sealed container. Spilled particles have to be returned to Siemens. | in an environmentally | 4.22% | 51 | 1123 | Make note in the relevant service documents. How these parts are disposed in an environmentally friendly manner has to be documented in the disposal instructions. Cause: Toxic materials Possible consequences: Toxic contamination in case of detector panel is broken by mechanical shock Remedy: Handle detector with care! If the detector is broken do not use it, inspect if any particles are spilled. Leaked particles have to be collected and kept in a sealed container. Spilled particles have to be returned to Siemens. | FALSE
The following general principles for designing the protective path shall be considered: - The software of the protective path shall be developed according to software safety class C (IEC 62304). - Common causes (errors or environmental conditions that could affect C and P simultaneously) shall be considered and excluded by design: ESD / transients (sufficient electrical isolation between P and C path signals), temperature out of specified limits, etc. | following general | 3.69% | 280 | 806 | The following general principles for designing the protective path shall be considered: - The software of the protective path shall be developed according to software safety class C (IEC 62304). - Common causes (errors or environmental conditions that could affect C and P simultaneously) shall be considered and excluded by design: ESD / transients (sufficient electrical isolation between P and C path signals), temperature out of specified limits, etc. | FALSE
CAUTION Cause: Storage rate is set incorrectly or being not aware of replay / image saving behavior Possible consequences: Images are lost or replay rate / image saving differs from the pulse rate during acquisition. Remedy: - Choose the storage rate with care. - Note that for 30 frames per second with 30x30 detector in MAG1 or with 20x20 detector in overview, image information is summarized in 15 frames for replay / storage only (no image information is lost). | image information | 4.03% | 26 | 649 | CAUTION Cause: Storage rate is set incorrectly or being not aware of replay / image saving behavior Possible consequences: Images are lost or replay rate / image saving differs from the pulse rate during acquisition. Remedy: - Choose the storage rate with care. - Note that for 30 frames per second with 30x30 detector in MAG1 or with 20x20 detector in overview, image information is summarized in 15 frames for replay / storage only (no image information is lost). | FALSE
The following time restrictions shall be met for recovery management (this is also valid for EMC immunity testing): Time from initiating a regular shut down via off button, booting the system again, until X-ray ready: 4 min Time from power supply interruption, booting the system again, until X-ray ready: 120 s Note: Due to the limited time to recover all system functions, no emergency radiography mode must be provided. | ecovery management | 4.2% | 268 | 616 | The following time restrictions shall be met for recovery management (this is also valid for EMC immunity testing): Time from initiating a regular shut down via off button, booting the system again, until X-ray ready: 4 min Time from power supply interruption, booting the system again, until X-ray ready: 120 s Note: Due to the limited time to recover all system functions, no emergency radiography mode must be provided. | FALSE
Service Safety Instruction WARNING Cause: If the ESU is not discharged completely, electrical voltage is present and is accessible when e. g. main unit covers are removed. Possible consequences: Risk of death or serious injury due to electric shock when electrified components are touched. Remedy: Wait until discharge finishes (10min if without ESU and 20min if equipped with ESU) for both ESU and generator capacitors before starting to work on relevant parts of the main unit. Observe the LED at the inverter box which is lit when the capacitors are charged and turned off nealy discharged. Don't disconnect ESU and generator before this time as well, otherwise the discharging circuit will not work. | death or serious injury | 3.37% | 4 | 108 | Service Safety Instruction WARNING Cause: If the ESU is not discharged completely, electrical voltage is present and is accessible when e. g. main unit covers are removed. Possible consequences: Risk of death or serious injury due to electric shock when electrified components are touched. Remedy: Wait until discharge finishes (10min if without ESU and 20min if equipped with ESU) for both ESU and generator capacitors before starting to work on relevant parts of the main unit. Observe the LED at the inverter box which is lit when the capacitors are charged and turned off nealy discharged. Don't disconnect ESU and generator before this time as well, otherwise the discharging circuit will not work. | FALSE
Warning Electrical voltage on filter Z1 There is no protection cover on connector of filter Z1. Live voltage is present at the power filter Z1. Even after the system switches off the power supply, electrical voltage is present.Live lines can cause electric shock Risk of death, serious physical injury, or damage to property. ->De-energize the system ->If checks or measurements require working with a live line, cover energized parts to prevent unintentional contact. ->Use extreme caution when working on live lines. ->Use only measurement equipment and devices that are suited to the task ->Read and observe the warning labels attached next to conducting parts. | fter the system switches off | 4.17% | 100 | 110 | Warning Electrical voltage on filter Z1 There is no protection cover on connector of filter Z1. Live voltage is present at the power filter Z1. Even after the system switches off the power supply, electrical voltage is present.Live lines can cause electric shock Risk of death, serious physical injury, or damage to property. ->De-energize the system ->If checks or measurements require working with a live line, cover energized parts to prevent unintentional contact. ->Use extreme caution when working on live lines. ->Use only measurement equipment and devices that are suited to the task ->Read and observe the warning labels attached next to conducting parts. | FALSE
Note: System specific documents (e.g. protocols etc., i.e. blue 'Service-Folder') may be stored inside of trolley. This wish come up as system specific documents needs to be on-site during product life time. Result of risk evaluation: temperature inside of trolley will never reach such threshold to enable auto-ignition or any melting of material. | temperature i | 3.98% | 328 | 1215 | Note: System specific documents (e.g. protocols etc., i.e. blue 'Service-Folder') may be stored inside of trolley. This wish come up as system specific documents needs to be on-site during product life time. Result of risk evaluation: temperature inside of trolley will never reach such threshold to enable auto-ignition or any melting of material. | FALSE
The chassis must be designed in a way that prevents damage of components when hitting a door frame. This is in accordance to IEC60601-1, clause 15.3.5 (Rough Handling Test). Door frame test, scenario 1: Scenario: hitting a door frame while moving Main Unit into OR (or from one OR to another) Setup: transport position is a prerequisite Speed: 0.8 m/s (tolerance: +/- 0.1 m/s) Quantity: repeat three times Door frame test, scenario 2: Scenario: hitting another system or wall while moving Main Unit within OR Setup: transport position is not a prerequisite Speed: 0.5 m/s (tolerance: +/- 0.1 m/s) Quantity: repeat ten times Remark: Execute door frame shock test according to standard, possibly emerged damages have to be examined for hazard (e.g. access to electrical and moveable parts has to be prevented). Door frame test, scenario 1: Scenario: hitting a door frame while moving Main Unit into OR (or from one OR to another) Setup: transport position is a prerequisite Speed: 0.8 m/s (tolerance: +/- 0.1 m/s) Quantity: repeat three times Door frame test, scenario 2: Scenario: hitting another system or wall while moving Main Unit within OR Setup: transport position is not a prerequisite Speed: 0.5 m/s (tolerance: +/- 0.1 m/s) Quantity: repeat ten times Remark: Execute door frame shock test according to standard, possibly emerged damages have to be examined for hazard (e.g. access to electrical and moveable parts has to be prevented). | in accordance to IEC | 1.43% | 266 | 1137 | The chassis must be designed in a way that prevents damage of components when hitting a door frame. This is in accordance to IEC60601-1, clause 15.3.5 (Rough Handling Test). Door frame test, scenario 1: Scenario: hitting a door frame while moving Main Unit into OR (or from one OR to another) Setup: transport position is a prerequisite Speed: 0.8 m/s (tolerance: +/- 0.1 m/s) Quantity: repeat three times Door frame test, scenario 2: Scenario: hitting another system or wall while moving Main Unit within OR Setup: transport position is not a prerequisite Speed: 0.5 m/s (tolerance: +/- 0.1 m/s) Quantity: repeat ten times Remark: Execute door frame shock test according to standard, possibly emerged damages have to be examined for hazard (e.g. access to electrical and moveable parts has to be prevented). Door frame test, scenario 1: Scenario: hitting a door frame while moving Main Unit into OR (or from one OR to another) Setup: transport position is a prerequisite Speed: 0.8 m/s (tolerance: +/- 0.1 m/s) Quantity: repeat three times Door frame test, scenario 2: Scenario: hitting another system or wall while moving Main Unit within OR Setup: transport position is not a prerequisite Speed: 0.5 m/s (tolerance: +/- 0.1 m/s) Quantity: repeat ten times Remark: Execute door frame shock test according to standard, possibly emerged damages have to be examined for hazard (e.g. access to electrical and moveable parts has to be prevented). | FALSE
WARNING Cause: Service, disassembly or refurbishment activities on hardware parts of Cios systems which had been in clinical use and are potentially contaminated. Possible consequences: Increased risk of infection due to residues of infectious substances such as bodily fluids or deposits of organic matter colonized with microorganisms Remedy: Consequently apply precautionary measures for handling potentially infectious material, especially: - Wear protective gloves, closed and long-sleeved protective clothing, protective shoes, and protective glasses. - In case of severe contamination or strong infectiousness, in case the contamination may be dusting or emit droplets, apply additional measures such as face masks and disposable protective clothing, for example overalls and caps. - After finishing the activities on contaminated hardware parts, clean and disinfect potentially contaminated body parts, especially hands and face, as well as clothing and equipment. | clean and disinfect | 2.14% | 257 | 194 | WARNING Cause: Service, disassembly or refurbishment activities on hardware parts of Cios systems which had been in clinical use and are potentially contaminated. Possible consequences: Increased risk of infection due to residues of infectious substances such as bodily fluids or deposits of organic matter colonized with microorganisms Remedy: Consequently apply precautionary measures for handling potentially infectious material, especially: - Wear protective gloves, closed and long-sleeved protective clothing, protective shoes, and protective glasses. - In case of severe contamination or strong infectiousness, in case the contamination may be dusting or emit droplets, apply additional measures such as face masks and disposable protective clothing, for example overalls and caps. - After finishing the activities on contaminated hardware parts, clean and disinfect potentially contaminated body parts, especially hands and face, as well as clothing and equipment. | FALSE
Meet IEC 60601-1-2 standards to avoid electromagnetic disturbances. 1. Electromagnetic Emissions: Conducted emissions as Class A Limits CISPR 11 Radiated emissions as Class A Limits CISPR 11 Regarding static magnetic fields (e. g. drives, open brakes, permanent magnets): The requirements of EU directive 2013/35/EU shall be fulfilled: max. 0,5 mT. This threshold shall be met in clinical use (threshold met outside of covers, at all parts where the operator or patient have access) and during service (threshold met at the surface of the individual component). If this is not possible in all scenarios by design, additional measures shall be defined (labels etc.). 2. Electromagnetic Immunity: General: The test levels for professional healthcare facility environment shall be applied. Electrostatic discharge according basic standard IEC 61000-4-2 (+/-8kV contact, +/-15kV air) Radiated RF electromagnetic fields according basic standard IEC 61000-4-3 (80Mhz to 2,5Ghz - 30V/m) Electrical fast transients and bursts according basic standard IEC 61000-4-4 Surges according basic standard IEC 61000-4-5 Conducted disturbances induced by RF fields according basic standard IEC 61000-4-6 3. Immunity with activated HF surgical equipment: The immunity shall be maintained in combination with activated HF surgical equipment. The CIOS systems do not provide a table, but nevertheless, HF surgical equipment is commonly used in the vicinity of the C-arm. Therefore, acceptance criteria for testing with HF surgical equipment are derived from a usability based approach, however in alignment with IEC 60601-2-2 and IEC 60601-2-46 as far as applicable. 4. Pass/fail criteria for disturbances (general system functions) The system shall be immune to the above mentioned test conditions. That means that the following disturbances or failures are not acceptable (immunity pass/fail criteria): - Unintended release of X-ray - Unclear radiation status (X-ray indicator lights up) - Unintended release of motorized movements - Loss of basic safety (e. g. electrical safety: insulation break during ESD testing) - Loss of essential performance required by particular standards, see Essential Performance, pg. 40 [273657] (note that dose report is part of patient data export) - Component failures (observe also error messages during testing) - Changes in programmable parameters - Reset to factory defaults - Change of operator mode - False alarms - Cessation or interruption of any intended operation (non-recoverable system failure) - Error of a displayed numerical value sufficiently large to affect diagnosis or treatment The following are examples for acceptable disturbances: - Flickering or expiration of LEDs - Activation or deactivation of the lasers - Audible interference of the sound system - Temporary loss of functionality on external interfaces (incl. NaviLink interface, InstantLink (NXS, NGS) interface) - Self-recoverable fail of single touchscreen panels (after end of disturbance) without faulty activation < 30 s (note, redundant touch entries available) 5. Pass/fail criteria for image display A) Fail: - Artifact or distortion in an image in which the artifact would interfere with diagnosis or treatment (e. g. not clearly distinguishable from clinical image content) B) Acceptable: - Minor image distortions or major short-term image distortions (if clearly distinguishable from the clinical image content) - If image artifacts occur, they shall be clearly distinguishable from the clinical image content - Flickering or self-recoverable display loss (after end of disturbance) on the main monitor < 1 s (note, display is not redundant) - Flickering or self-recoverable display loss (after end of disturbance) on any operator panel < 5 s (note, the operator panel displays are redundant) Note: Subsystem AppHost only suitable for system with it. | for professional healthcare facility environment | 1.25% | 271 | 1207 | Meet IEC 60601-1-2 standards to avoid electromagnetic disturbances. 1. Electromagnetic Emissions: Conducted emissions as Class A Limits CISPR 11 Radiated emissions as Class A Limits CISPR 11 Regarding static magnetic fields (e. g. drives, open brakes, permanent magnets): The requirements of EU directive 2013/35/EU shall be fulfilled: max. 0,5 mT. This threshold shall be met in clinical use (threshold met outside of covers, at all parts where the operator or patient have access) and during service (threshold met at the surface of the individual component). If this is not possible in all scenarios by design, additional measures shall be defined (labels etc.). 2. Electromagnetic Immunity: General: The test levels for professional healthcare facility environment shall be applied. Electrostatic discharge according basic standard IEC 61000-4-2 (+/-8kV contact, +/-15kV air) Radiated RF electromagnetic fields according basic standard IEC 61000-4-3 (80Mhz to 2,5Ghz - 30V/m) Electrical fast transients and bursts according basic standard IEC 61000-4-4 Surges according basic standard IEC 61000-4-5 Conducted disturbances induced by RF fields according basic standard IEC 61000-4-6 3. Immunity with activated HF surgical equipment: The immunity shall be maintained in combination with activated HF surgical equipment. The CIOS systems do not provide a table, but nevertheless, HF surgical equipment is commonly used in the vicinity of the C-arm. Therefore, acceptance criteria for testing with HF surgical equipment are derived from a usability based approach, however in alignment with IEC 60601-2-2 and IEC 60601-2-46 as far as applicable. 4. Pass/fail criteria for disturbances (general system functions) The system shall be immune to the above mentioned test conditions. That means that the following disturbances or failures are not acceptable (immunity pass/fail criteria): - Unintended release of X-ray - Unclear radiation status (X-ray indicator lights up) - Unintended release of motorized movements - Loss of basic safety (e. g. electrical safety: insulation break during ESD testing) - Loss of essential performance required by particular standards, see Essential Performance, pg. 40 [273657] (note that dose report is part of patient data export) - Component failures (observe also error messages during testing) - Changes in programmable parameters - Reset to factory defaults - Change of operator mode - False alarms - Cessation or interruption of any intended operation (non-recoverable system failure) - Error of a displayed numerical value sufficiently large to affect diagnosis or treatment The following are examples for acceptable disturbances: - Flickering or expiration of LEDs - Activation or deactivation of the lasers - Audible interference of the sound system - Temporary loss of functionality on external interfaces (incl. NaviLink interface, InstantLink (NXS, NGS) interface) - Self-recoverable fail of single touchscreen panels (after end of disturbance) without faulty activation < 30 s (note, redundant touch entries available) 5. Pass/fail criteria for image display A) Fail: - Artifact or distortion in an image in which the artifact would interfere with diagnosis or treatment (e. g. not clearly distinguishable from clinical image content) B) Acceptable: - Minor image distortions or major short-term image distortions (if clearly distinguishable from the clinical image content) - If image artifacts occur, they shall be clearly distinguishable from the clinical image content - Flickering or self-recoverable display loss (after end of disturbance) on the main monitor < 1 s (note, display is not redundant) - Flickering or self-recoverable display loss (after end of disturbance) on any operator panel < 5 s (note, the operator panel displays are redundant) Note: Subsystem AppHost only suitable for system with it. | FALSE